## summary
Oxygen Calibration
AZMP Survey Teleost 200
June 27 - July 13, 2019
The following instruments were used:
3313: 708 records as primary and 0 records as secondary
3539: 0 records as primary and 708 records as secondary
The following stations were occupied:
BB: 01, 06, 10, 14
FC: 03, 06, 14, 20, 24, 33
MB: 04, 07, 11
SI: 04, 09, 13
The regression data is:
inst | slope | intercept | R2 | OldSOC | NewSOC
3313 | 1.0853 | 0.47903 | 0.96939 | 0.4339 | 0.5016
3539 | 1.1237 | 0.28414 | 0.9792 | 0.4637 | 0.5407

## data
Station-ID | Date | Time | ShTrpStn | Stickr | latD | latM | lonD | lonM | PrDM | PrDM_sdv | T090C | T090C_sdv | T190C | T190C_sdv | C0S/m | C0S/m_sdv | C1S/m | C1S/m_sdv | Sal00 | Sal00_sdv | Sal11 | Sal11_sdv | Sigma-t00 | Sigma-t00_sdv | Sigma-t11 | Sigma-t11_sdv | Sbeox0V | Sbeox0V_sdv | Sbeox1V | Sbeox1V_sdv | Sbeox0ML/L | Sbeox0ML/L_sdv | Sbeox1ML/L | Sbeox1ML/L_sdv | OxsolML/L | OxsolML/L_sdv | Sbeox0PS | Sbeox0PS_sdv | Sbeox1PS | Sbeox1PS_sdv | FlECO-AFL | FlECO-AFL_sdv | Par/lo | Par/lo_sdv | P | P_sdv | WetCDOM | CStarAt0 | CStarAt0_sdv | CStarTr | CStarTr_sdv | Analysis_Date_time | Sample | [O2](ml/l) | std_[O2](ml/l) | Rep_1(ml/l) | Rep_2(ml/l) | Filename | Instr | PrimOSens | PrimSOC | SecOSens | SecSOC | WINKLER | WINKLER-DIFF | section
FC-03 | 2019-06-27 | 09:13:03 | 39200005 | 433878 | 47 | 0 | -52 | 34.8 | 158.56 | 0.144 | -0.6905 | 0.0003 | -0.6906 | 0.0003 | 2.6993 | 2.6e-05 | 2.6994 | 3.9e-05 | 32.9422 | #NUM! | 32.9435 | #NUM! | 26.4746 | #NUM! | 26.4756 | #NUM! | 2.1371 | 0.0006 | 2.0086 | 0.0005 | 6.0142 | #NUM! | 5.9195 | #NUM! | 8.2758 | #NUM! | 72.672 | #NUM! | 71.527 | #NUM! | 0.1177 | 0.0053 | 0.11786 | 5.8571e-05 | 8.582 | 0 | 2.6158 | -0.0225 | 0.0009 | 100.5645 | 0.0226 | 2019-06-27 15:05:10 | 433878 | 6.717 | 0.042 | 6.747 | 6.688 | 39200005 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.7175 | 0.059 | FC
FC-03 | 2019-06-27 | 09:17:31 | 39200005 | 433879 | 47 | 0 | -52 | 34.8 | 100.879 | 0.186 | -0.6597 | 0.0061 | -0.6612 | 0.0062 | 2.6772 | 0.000416 | 2.6772 | 0.000427 | 32.6458 | #NUM! | 32.6475 | #NUM! | 26.2336 | #NUM! | 26.235 | #NUM! | 2.6271 | 0.0006 | 2.4546 | 0.0004 | 7.7662 | #NUM! | 7.6345 | #NUM! | 8.2864 | #NUM! | 93.723 | #NUM! | 92.132 | #NUM! | 0.0601 | 0.0057 | 0.11776 | 0.00018289 | 8.816 | 0.001 | 1.3829 | -0.0407 | 0.0002 | 101.0213 | 0.0054 | 2019-06-27 15:17:07 | 433879 | 8.695 | 0.047 | 8.662 | 8.729 | 39200005 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.6955 | 0.067 | FC
FC-03 | 2019-06-27 | 09:19:41 | 39200005 | 433880 | 47 | 0 | -52 | 34.8 | 75.583 | 0.113 | 0.0616 | 0.0056 | 0.0616 | 0.0102 | 2.7289 | 0.000482 | 2.729 | 0.000889 | 32.5749 | #NUM! | 32.576 | #NUM! | 26.1459 | #NUM! | 26.1468 | #NUM! | 2.7227 | 0.0009 | 2.5401 | 0.0006 | 7.9062 | #NUM! | 7.7682 | #NUM! | 8.1343 | #NUM! | 97.196 | #NUM! | 95.5 | #NUM! | 0.0987 | 0.0037 | 0.11769 | 0.00020604 | 8.869 | 0 | 0.8418 | -0.0386 | 0.0003 | 100.9709 | 0.0083 | 2019-06-27 15:29:57 | 433880 | 8.847 | 0.124 | 8.759 | 8.935 | 39200005 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.847 | 0.176 | FC
FC-03 | 2019-06-27 | 09:21:58 | 39200005 | 433881 | 47 | 0 | -52 | 34.8 | 51.008 | 0.027 | 0.6067 | 0.0009 | 0.6096 | 0.0008 | 2.7676 | 9.2e-05 | 2.7679 | 6.5e-05 | 32.5177 | #NUM! | 32.5186 | #NUM! | 26.072 | #NUM! | 26.0726 | #NUM! | 2.7572 | 0.0007 | 2.5748 | 0.0008 | 7.8875 | #NUM! | 7.7524 | #NUM! | 8.0228 | #NUM! | 98.315 | #NUM! | 96.63 | #NUM! | 1.0525 | 0.0133 | 0.93615 | 0.0079287 | 8.885 | 0 | 1.0414 | -0.0441 | 0.0004 | 101.1082 | 0.0101 | 2019-06-27 16:15:06 | 433881 | 8.803 | 0.008 | 8.798 | 8.809 | 39200005 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.8035 | 0.011 | FC
FC-03 | 2019-06-27 | 09:23:34 | 39200005 | 433882 | 47 | 0 | -52 | 34.8 | 41.009 | 0.188 | 0.8195 | 0.1084 | 0.8069 | 0.0319 | 2.7786 | 0.008347 | 2.7782 | 0.00266 | 32.4423 | #NUM! | 32.4492 | #NUM! | 25.9995 | #NUM! | 26.0058 | #NUM! | 2.7624 | 0.002 | 2.5865 | 0.004 | 7.8383 | #NUM! | 7.644 | #NUM! | 7.9829 | #NUM! | 98.19 | #NUM! | 95.756 | #NUM! | 0.6635 | 0.0104 | 2.6724 | 0.063131 | 8.878 | 0.001 | 1.6066 | -0.0339 | 0.0009 | 100.8508 | 0.0231 | 2019-06-27 16:26:32 | 433882 | 8.726 | 0.006 | 8.721 | 8.73 | 39200005 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.7255 | 0.009 | FC
FC-03 | 2019-06-27 | 09:25:23 | 39200005 | 433883 | 47 | 0 | -52 | 34.8 | 31.458 | 0.12 | 2.4129 | 0.0011 | 2.4116 | 0.0003 | 2.9056 | 8.7e-05 | 2.9056 | 2.2e-05 | 32.4027 | #NUM! | 32.4039 | #NUM! | 25.8601 | #NUM! | 25.8611 | #NUM! | 2.799 | 0.0005 | 2.6141 | 0.0006 | 7.6007 | #NUM! | 7.4513 | #NUM! | 7.6677 | #NUM! | 99.127 | #NUM! | 97.179 | #NUM! | 0.7836 | 0.0051 | 6.6168 | 0.13151 | 8.867 | 0 | 1.3955 | -0.0402 | 0.0006 | 101.0107 | 0.0136 | 2019-06-27 16:37:00 | 433883 | 8.422 | 0.021 | 8.407 | 8.436 | 39200005 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.4215 | 0.029 | FC
FC-03 | 2019-06-27 | 09:27:09 | 39200005 | 433884 | 47 | 0 | -52 | 34.8 | 19.992 | 0.051 | 4.2782 | 0.0022 | 4.2809 | 0.0022 | 3.0507 | 0.000166 | 3.051 | 0.00017 | 32.2861 | #NUM! | 32.2869 | #NUM! | 25.6002 | #NUM! | 25.6006 | #NUM! | 2.8431 | 0.0004 | 2.6565 | 0.0009 | 7.3507 | #NUM! | 7.1915 | #NUM! | 7.3291 | #NUM! | 100.298 | #NUM! | 98.125 | #NUM! | 0.7712 | 0.0059 | 22.301 | 0.092823 | 8.86 | 0 | 1.2058 | 0.0364 | 0.0004 | 99.0945 | 0.0102 | 2019-06-27 16:47:26 | 433884 | 8.173 | 0.003 | 8.175 | 8.171 | 39200005 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.173 | 0.004 | FC
FC-03 | 2019-06-27 | 09:28:41 | 39200005 | 433885 | 47 | 0 | -52 | 34.8 | 10.026 | 0.045 | 4.9958 | 0.0031 | 4.9983 | 0.0036 | 3.1065 | 0.000196 | 3.1068 | 0.000256 | 32.2386 | #NUM! | 32.2393 | #NUM! | 25.487 | #NUM! | 25.4872 | #NUM! | 2.8466 | 0.001 | 2.665 | 0.0008 | 7.2202 | #NUM! | 7.0756 | #NUM! | 7.2059 | #NUM! | 100.202 | #NUM! | 98.195 | #NUM! | 0.5528 | 0.0037 | 72.822 | 4.7055 | 8.86 | 0.001 | 0.9295 | 0.0304 | 0.0079 | 99.2424 | 0.1948 | 2019-06-27 16:57:20 | 433885 | 8.063 | 0.047 | 8.097 | 8.03 | 39200005 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.0635 | 0.067 | FC
FC-03 | 2019-06-27 | 09:30:09 | 39200005 | 433886 | 47 | 0 | -52 | 34.8 | 5.267 | 0.157 | 5.6581 | 0.0204 | 5.6587 | 0.0254 | 3.162 | 0.001829 | 3.1621 | 0.002111 | 32.2345 | #NUM! | 32.2359 | #NUM! | 25.4087 | #NUM! | 25.4097 | #NUM! | 2.8603 | 0.0008 | 2.672 | 0.0004 | 7.1174 | #NUM! | 6.976 | #NUM! | 7.0936 | #NUM! | 100.339 | #NUM! | 98.345 | #NUM! | 0.3432 | 0.005 | 162.4 | 31.224 | 8.85 | 0 | 2.0764 | 0.0112 | 0.0008 | 99.7202 | 0.0192 | 2019-06-27 17:06:51 | 433886 | 7.962 | 0.011 | 7.969 | 7.954 | 39200005 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.9615 | 0.015 | FC
FC-06 | 2019-06-27 | 17:57:57 | 39200008 | 433907 | 47 | 0 | -51 | 29.1 | 100.077 | 0.112 | -0.5467 | 0.0005 | -0.5468 | 0.0009 | 2.698 | 2.2e-05 | 2.6981 | 2.6e-05 | 32.803 | #NUM! | 32.8044 | #NUM! | 26.3564 | #NUM! | 26.3576 | #NUM! | 2.0039 | 0.0007 | 1.8779 | 0.0008 | 5.4668 | #NUM! | 5.3645 | #NUM! | 8.2524 | #NUM! | 66.244 | #NUM! | 65.005 | #NUM! | 0.2131 | 0.0056 | 0.11771 | 0.00020184 | 8.455 | 0 | 1.0693 | 0.0416 | 0.0016 | 98.9644 | 0.0399 | 2019-06-27 22:51:44 | 433907 | 6.225 | 0.011 | 6.233 | 6.217 | 39200008 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.225 | 0.016 | FC
FC-06 | 2019-06-27 | 18:01:27 | 39200008 | 433908 | 47 | 0 | -51 | 29.1 | 77.925 | 0.045 | 0.2072 | 0.0038 | 0.2067 | 0.0033 | 2.7435 | 0.000262 | 2.7436 | 0.000232 | 32.6102 | #NUM! | 32.6117 | #NUM! | 26.1673 | #NUM! | 26.1685 | #NUM! | 2.45 | 0.0013 | 2.2798 | 0.0008 | 6.933 | #NUM! | 6.7547 | #NUM! | 8.1014 | #NUM! | 85.578 | #NUM! | 83.378 | #NUM! | 0.0919 | 0.0086 | 0.40472 | 0.019826 | 8.692 | 0.002 | 1.2656 | -0.0778 | 0.0005 | 101.9653 | 0.0118 | 2019-06-27 23:24:23 | 433908 | 7.806 | 0.033 | 7.783 | 7.829 | 39200008 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.806 | 0.046 | FC
FC-06 | 2019-06-27 | 18:04:11 | 39200008 | 433909 | 47 | 0 | -51 | 29.1 | 50.758 | 0.112 | 1.019 | 0.0011 | 1.0193 | 0.0012 | 2.8018 | 9e-05 | 2.8019 | 9e-05 | 32.5251 | #NUM! | 32.526 | #NUM! | 26.0544 | #NUM! | 26.0551 | #NUM! | 2.7432 | 0.0006 | 2.5464 | 0.0006 | 7.742 | #NUM! | 7.5413 | #NUM! | 7.9374 | #NUM! | 97.539 | #NUM! | 95.011 | #NUM! | 0.1339 | 0.0048 | 2.9583 | 0.02982 | 8.867 | 0.001 | 2.493 | -0.0684 | 0.0003 | 101.7244 | 0.009 | 2019-06-27 23:36:31 | 433909 | 8.605 | 0.066 | 8.652 | 8.558 | 39200008 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.605 | 0.094 | FC
FC-06 | 2019-06-27 | 18:05:51 | 39200008 | 433910 | 47 | 0 | -51 | 29.1 | 40.145 | 0.045 | 1.1831 | 0.0061 | 1.1804 | 0.0032 | 2.8148 | 0.000528 | 2.8146 | 0.000275 | 32.5258 | #NUM! | 32.5267 | #NUM! | 26.0449 | #NUM! | 26.0458 | #NUM! | 2.7842 | 0.0005 | 2.5848 | 0.0006 | 7.8336 | #NUM! | 7.6362 | #NUM! | 7.9039 | #NUM! | 99.111 | #NUM! | 96.613 | #NUM! | 1.1217 | 0.025 | 7.2268 | 0.025543 | 8.881 | 0 | 1.4026 | -0.0576 | 0.0044 | 101.4492 | 0.1107 | 2019-06-27 23:47:15 | 433910 | 8.695 | 0.017 | 8.683 | 8.707 | 39200008 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.695 | 0.024 | FC
FC-06 | 2019-06-27 | 18:08:06 | 39200008 | 433911 | 47 | 0 | -51 | 29.1 | 29.588 | 0.088 | 3.2309 | 0.0321 | 3.2317 | 0.0315 | 2.9798 | 0.002525 | 2.9799 | 0.002466 | 32.4803 | #NUM! | 32.4802 | #NUM! | 25.854 | #NUM! | 25.8538 | #NUM! | 2.813 | 0.0015 | 2.615 | 0.0017 | 7.4842 | #NUM! | 7.3063 | #NUM! | 7.5093 | #NUM! | 99.667 | #NUM! | 97.299 | #NUM! | 0.553 | 0.0095 | 19.141 | 0.1833 | 8.865 | 0.002 | 1.2201 | -0.026 | 0.0043 | 100.6508 | 0.1087 | 2019-06-27 23:59:11 | 433911 | 8.149 | 0.012 | 8.157 | 8.14 | 39200008 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.1485 | 0.017 | FC
FC-06 | 2019-06-27 | 18:10:28 | 39200008 | 433912 | 47 | 0 | -51 | 29.1 | 19.87 | 0.103 | 4.8639 | 0.0038 | 4.8642 | 0.0038 | 3.1115 | 0.000331 | 3.1116 | 0.000338 | 32.4192 | #NUM! | 32.4207 | #NUM! | 25.6445 | #NUM! | 25.6456 | #NUM! | 2.827 | 0.0003 | 2.627 | 0.0009 | 7.1775 | #NUM! | 6.9723 | #NUM! | 7.2199 | #NUM! | 99.415 | #NUM! | 96.573 | #NUM! | 0.2787 | 0.0076 | 35.776 | 0.28569 | 8.865 | 0 | 0.9471 | 0.0013 | 0.0014 | 99.9665 | 0.0345 | 2019-06-28 00:09:41 | 433912 | 7.956 | 0.004 | 7.959 | 7.953 | 39200008 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.956 | 0.006 | FC
FC-06 | 2019-06-27 | 18:12:58 | 39200008 | 433913 | 47 | 0 | -51 | 29.1 | 9.985 | 0.168 | 6.9823 | 0.0075 | 6.9777 | 0.0127 | 3.2914 | 0.00063 | 3.2912 | 0.001019 | 32.4116 | #NUM! | 32.413 | #NUM! | 25.3826 | #NUM! | 25.3844 | #NUM! | 2.8857 | 0.001 | 2.6809 | 0.0006 | 6.9616 | #NUM! | 6.7546 | #NUM! | 6.8698 | #NUM! | 101.341 | #NUM! | 98.326 | #NUM! | 0.2904 | 0.0052 | 75.886 | 1.2484 | 8.84 | 0.003 | 1.1198 | 0.0103 | 0.0004 | 99.7418 | 0.011 | 2019-06-27 23:08:34 | 433913 | 7.78 | 0.006 | 7.775 | 7.784 | 39200008 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.7795 | 0.009 | FC
FC-06 | 2019-06-27 | 18:14:30 | 39200008 | 433914 | 47 | 0 | -51 | 29.1 | 6.069 | 0.085 | 7.1234 | 0.0031 | 7.1259 | 0.0022 | 3.3036 | 0.000242 | 3.304 | 0.0002 | 32.4135 | #NUM! | 32.4151 | #NUM! | 25.3653 | #NUM! | 25.3662 | #NUM! | 2.8914 | 0.0006 | 2.6914 | 0 | 6.9382 | #NUM! | 6.7656 | #NUM! | 6.8475 | #NUM! | 101.329 | #NUM! | 98.809 | #NUM! | 0.2841 | 0.0048 | 111.84 | 3.4523 | 8.837 | 0.002 | 2.1652 | 0.0095 | 0.0005 | 99.7628 | 0.0136 | 2019-06-27 23:02:41 | 433914 | 7.769 | 0 | 7.769 | #NUM! | 39200008 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.769 | #NUM! | FC
FC-14 | 2019-06-28 | 17:06:39 | 39200016 | 433958 | 47 | 0 | -47 | 30 | 211.335 | 0.139 | 2.1286 | 0.0005 | 2.1294 | 0.0005 | 3.0379 | 0.000426 | 3.0381 | 5.4e-05 | 34.2396 | #NUM! | 34.2416 | #NUM! | 27.3523 | #NUM! | 27.3539 | #NUM! | 2.1247 | 0.0005 | 2.0056 | 0.0006 | 5.4774 | #NUM! | 5.4198 | #NUM! | 7.6259 | #NUM! | 71.827 | #NUM! | 71.072 | #NUM! | 0.0756 | 0.0041 | 0.11764 | 0.00020604 | 8.648 | 0.001 | -0.0595 | 0.0977 | 0.0051 | 97.5886 | 0.1243 | 2019-06-28 22:07:50 | 433958 | 6.16 | 0.029 | 6.181 | 6.14 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.1605 | 0.041 | FC
FC-14 | 2019-06-28 | 17:09:48 | 39200016 | 433959 | 47 | 0 | -47 | 30 | 150.369 | 0.039 | -0.5026 | 0.0003 | -0.5014 | 0.0002 | 2.7332 | 2.5e-05 | 2.7333 | 2.8e-05 | 33.1964 | #NUM! | 33.1972 | #NUM! | 26.6728 | #NUM! | 26.6734 | #NUM! | 2.3224 | 0.0005 | 2.1884 | 0.0002 | 6.6474 | #NUM! | 6.5759 | #NUM! | 8.22 | #NUM! | 80.868 | #NUM! | 79.999 | #NUM! | 0.0487 | 0.0034 | 0.11774 | 0.00019426 | 8.713 | 0.002 | 2.4963 | -0.0756 | 0.0005 | 101.9071 | 0.0133 | 2019-06-28 22:19:36 | 433959 | 7.525 | 0.001 | 7.526 | 7.524 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.525 | 0.002 | FC
FC-14 | 2019-06-28 | 17:12:22 | 39200016 | 433960 | 47 | 0 | -47 | 30 | 100.746 | 0.06 | -0.9984 | 0.0039 | -1.0053 | 0.0119 | 2.6672 | 0.000275 | 2.6668 | 0.000793 | 32.8797 | #NUM! | 32.881 | #NUM! | 26.4349 | #NUM! | 26.4362 | #NUM! | 2.5315 | 0.0005 | 2.3745 | 0.0005 | 7.4862 | #NUM! | 7.388 | #NUM! | 8.3476 | #NUM! | 89.68 | #NUM! | 88.503 | #NUM! | 0.0307 | 0.0061 | 0.24061 | 0.0051559 | 8.808 | 0 | 1.7146 | -0.0562 | 0.0006 | 101.4144 | 0.0136 | 2019-06-28 22:30:35 | 433960 | 8.352 | 0.001 | 8.353 | 8.351 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.352 | 0.002 | FC
FC-14 | 2019-06-28 | 17:14:13 | 39200016 | 433961 | 47 | 0 | -47 | 30 | 75.146 | 0.032 | -0.8582 | 0.0002 | -0.8576 | 0.0004 | 2.6724 | 1.8e-05 | 2.6724 | 2.4e-05 | 32.8117 | #NUM! | 32.8121 | #NUM! | 26.375 | #NUM! | 26.3754 | #NUM! | 2.5729 | 0.0006 | 2.4146 | 0.0002 | 7.5877 | #NUM! | 7.4887 | #NUM! | 8.3205 | #NUM! | 91.192 | #NUM! | 90.003 | #NUM! | 0.0605 | 0.0236 | 1.3687 | 0.0076028 | 8.827 | 0 | 1.816 | -0.0471 | 0.0001 | 101.1845 | 0.0039 | 2019-06-28 22:40:54 | 433961 | 8.465 | 0.007 | 8.46 | 8.47 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.465 | 0.01 | FC
FC-14 | 2019-06-28 | 17:15:52 | 39200016 | 433962 | 47 | 0 | -47 | 30 | 50.291 | 0.045 | -0.2329 | 0.0011 | -0.2328 | 0.0013 | 2.7162 | 9.6e-05 | 2.7163 | 0.000128 | 32.7375 | #NUM! | 32.7379 | #NUM! | 26.2904 | #NUM! | 26.2908 | #NUM! | 2.6987 | 0.0006 | 2.5275 | 0.0007 | 7.8715 | #NUM! | 7.75 | #NUM! | 8.1881 | #NUM! | 96.133 | #NUM! | 94.649 | #NUM! | 0.0415 | 0.0043 | 6.2767 | 0.028595 | 8.885 | 0.001 | 2.4129 | -0.0438 | 0.0003 | 101.101 | 0.0083 | 2019-06-28 22:52:07 | 433962 | 8.708 | 0.013 | 8.717 | 8.699 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.708 | 0.018 | FC
FC-14 | 2019-06-28 | 17:17:12 | 39200016 | 433963 | 47 | 0 | -47 | 30 | 40.27 | 0.035 | 0.1665 | 0.0004 | 0.1689 | 0.0003 | 2.7439 | 3e-05 | 2.7441 | 2.6e-05 | 32.6808 | #NUM! | 32.681 | #NUM! | 26.2262 | #NUM! | 26.2263 | #NUM! | 2.7491 | 0.0006 | 2.5781 | 0.0003 | 7.9389 | #NUM! | 7.8378 | #NUM! | 8.106 | #NUM! | 97.939 | #NUM! | 96.691 | #NUM! | 0.1086 | 0.0177 | 11.184 | 0.046472 | 8.903 | 0.002 | 2.5303 | -0.0439 | 0.0002 | 101.1044 | 0.0039 | 2019-06-28 23:03:31 | 433963 | 8.873 | 0.009 | 8.88 | 8.867 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.8735 | 0.013 | FC
FC-14 | 2019-06-28 | 17:18:32 | 39200016 | 433964 | 47 | 0 | -47 | 30 | 30.308 | 0.152 | 1.9069 | 0.0161 | 1.8933 | 0.0193 | 2.8791 | 0.001183 | 2.8782 | 0.001469 | 32.5917 | #NUM! | 32.5948 | #NUM! | 26.0493 | #NUM! | 26.0527 | #NUM! | 2.7796 | 0.0006 | 2.6081 | 0.0007 | 7.6412 | #NUM! | 7.543 | #NUM! | 7.756 | #NUM! | 98.521 | #NUM! | 97.254 | #NUM! | 0.3861 | 0.0071 | 17.797 | 0.29887 | 8.933 | 0.001 | 0.8314 | -0.0508 | 0.0006 | 101.2787 | 0.0137 | 2019-06-28 23:14:48 | 433964 | 8.591 | 0.03 | 8.612 | 8.57 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.591 | 0.042 | FC
FC-14 | 2019-06-28 | 17:20:05 | 39200016 | 433965 | 47 | 0 | -47 | 30 | 20.592 | 0.038 | 3.582 | 0.0035 | 3.5911 | 0.0028 | 3.0133 | 0.000277 | 3.0141 | 0.000226 | 32.5346 | #NUM! | 32.534 | #NUM! | 25.8653 | #NUM! | 25.864 | #NUM! | 2.7884 | 0.0005 | 2.6176 | 0.0008 | 7.3038 | #NUM! | 7.193 | #NUM! | 7.442 | #NUM! | 98.146 | #NUM! | 96.657 | #NUM! | 1.5209 | 0.0205 | 39.891 | 0.35095 | 8.917 | 0 | 2.095 | 0.0787 | 0.0012 | 98.0514 | 0.0288 | 2019-06-28 23:26:02 | 433965 | 8.242 | 0.015 | 8.253 | 8.232 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.2425 | 0.021 | FC
FC-14 | 2019-06-28 | 17:21:42 | 39200016 | 433966 | 47 | 0 | -47 | 30 | 10.604 | 0.159 | 6.0742 | 0.0614 | 5.9613 | 0.0559 | 3.22 | 0.005212 | 3.2103 | 0.00454 | 32.4852 | #NUM! | 32.4843 | #NUM! | 25.5567 | #NUM! | 25.5698 | #NUM! | 2.8916 | 0.0027 | 2.7207 | 0.0041 | 7.1076 | #NUM! | 6.9911 | #NUM! | 7.0128 | #NUM! | 101.354 | #NUM! | 99.693 | #NUM! | 0.3054 | 0.0286 | 116.01 | 2.9608 | 8.886 | 0.001 | 0.6022 | 0.0395 | 0.001 | 99.0181 | 0.0245 | 2019-06-28 23:36:39 | 433966 | 8.029 | 0.018 | 8.042 | 8.017 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.0295 | 0.025 | FC
FC-14 | 2019-06-28 | 17:22:42 | 39200016 | 433967 | 47 | 0 | -47 | 30 | 4.179 | 0.18 | 5.9281 | 0.0486 | 5.9499 | 0.0568 | 3.2067 | 0.003916 | 3.2087 | 0.004647 | 32.4789 | #NUM! | 32.4807 | #NUM! | 25.5696 | #NUM! | 25.5684 | #NUM! | 2.8888 | 0.0014 | 2.7175 | 0.0006 | 7.1331 | #NUM! | 7.0583 | #NUM! | 7.0372 | #NUM! | 101.366 | #NUM! | 100.303 | #NUM! | 0.2301 | 0.0056 | 266.94 | 10.969 | 8.89 | 0.002 | 0.2064 | 0.0491 | 0.0014 | 98.78 | 0.0344 | 2019-06-28 23:47:57 | 433967 | 8.042 | 0.01 | 8.049 | 8.035 | 39200016 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.042 | 0.014 | FC
FC-20 | 2019-06-29 | 09:16:14 | 39200022 | 434001 | 47 | 0 | -46 | 29 | 351.655 | 0.158 | 4.0107 | 0.0013 | 4.0097 | 0.0018 | 3.2643 | 0.000113 | 3.2642 | 0.000153 | 34.9032 | #NUM! | 34.904 | #NUM! | 27.7081 | #NUM! | 27.7089 | #NUM! | 2.2215 | 0.0009 | 2.1161 | 0.0006 | 5.5841 | #NUM! | 5.5856 | #NUM! | 7.2476 | #NUM! | 77.049 | #NUM! | 77.07 | #NUM! | 0.0252 | 0.0042 | 0.11786 | 5.8571e-05 | 8.731 | 0.002 | 1.0468 | -0.0841 | 0.0005 | 102.1236 | 0.0135 | 2019-06-29 16:21:34 | 434001 | 6.499 | 0.011 | 6.507 | 6.491 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.499 | 0.016 | FC
FC-20 | 2019-06-29 | 09:23:08 | 39200022 | 434002 | 47 | 0 | -46 | 29 | 150.899 | 0.283 | 4.1209 | 0.0008 | 4.121 | 0.0006 | 3.2471 | 7.5e-05 | 3.2472 | 5.6e-05 | 34.6913 | #NUM! | 34.6921 | #NUM! | 27.5279 | #NUM! | 27.5286 | #NUM! | 2.2275 | 0.0008 | 2.12 | 0.0006 | 5.4378 | #NUM! | 5.4327 | #NUM! | 7.2386 | #NUM! | 75.123 | #NUM! | 75.052 | #NUM! | 0.0268 | 0.0082 | 0.11769 | 0.000205 | 8.732 | 0 | 0.1269 | -0.1105 | 0.0009 | 102.8002 | 0.0235 | 2019-06-29 16:06:18 | 434002 | 6.303 | 0.009 | 6.296 | 6.309 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.3025 | 0.013 | FC
FC-20 | 2019-06-29 | 09:25:43 | 39200022 | 434003 | 47 | 0 | -46 | 29 | 100.322 | 0.087 | 3.5895 | 0.0413 | 3.5833 | 0.0426 | 3.1741 | 0.004037 | 3.1736 | 0.004202 | 34.4085 | #NUM! | 34.4099 | #NUM! | 27.357 | #NUM! | 27.3588 | #NUM! | 2.2744 | 0.001 | 2.162 | 0.001 | 5.6432 | #NUM! | 5.6411 | #NUM! | 7.3467 | #NUM! | 76.815 | #NUM! | 76.787 | #NUM! | 0.0376 | 0.0056 | 0.11765 | 0.00020674 | 8.735 | 0.001 | 0.1451 | -0.1087 | 0.0014 | 102.7537 | 0.0358 | 2019-06-29 15:55:52 | 434003 | 6.629 | 0 | 6.629 | 6.629 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.629 | 0 | FC
FC-20 | 2019-06-29 | 09:27:43 | 39200022 | 434004 | 47 | 0 | -46 | 29 | 76.891 | 0.234 | 3.5623 | 0.0158 | 3.5609 | 0.0203 | 3.1555 | 0.001386 | 3.1555 | 0.001773 | 34.2271 | #NUM! | 34.2281 | #NUM! | 27.2152 | #NUM! | 27.2161 | #NUM! | 2.3263 | 0.0015 | 2.2112 | 0.0003 | 5.7916 | #NUM! | 5.7986 | #NUM! | 7.3606 | #NUM! | 78.685 | #NUM! | 78.78 | #NUM! | 0.0515 | 0.0096 | 0.11771 | 0.00020184 | 8.75 | 0 | 0.2158 | -0.1063 | 0.0005 | 102.6933 | 0.0135 | 2019-06-29 15:42:14 | 434004 | 6.739 | 0.004 | 6.736 | 6.742 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.739 | 0.006 | FC
FC-20 | 2019-06-29 | 09:29:43 | 39200022 | 434005 | 47 | 0 | -46 | 29 | 49.988 | 0.129 | 3.4058 | 0.0293 | 3.4117 | 0.024 | 3.1179 | 0.002972 | 3.1186 | 0.002365 | 33.9536 | #NUM! | 33.956 | #NUM! | 27.0123 | #NUM! | 27.0137 | #NUM! | 2.5778 | 0.0007 | 2.4411 | 0.0015 | 6.6308 | #NUM! | 6.5722 | #NUM! | 7.4026 | #NUM! | 89.576 | #NUM! | 88.784 | #NUM! | 0.7111 | 0.0289 | 0.96417 | 0.0090111 | 8.856 | 0 | -0.6126 | -0.0465 | 0.0019 | 101.1691 | 0.0476 | 2019-06-29 15:15:25 | 434005 | 7.666 | 0.028 | 7.686 | 7.646 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.666 | 0.04 | FC
FC-20 | 2019-06-29 | 09:31:43 | 39200022 | 434006 | 47 | 0 | -46 | 29 | 38.921 | 0.172 | 4.2592 | 0.0329 | 4.2551 | 0.0349 | 3.1631 | 0.00234 | 3.1628 | 0.0025 | 33.6146 | #NUM! | 33.616 | #NUM! | 26.6575 | #NUM! | 26.659 | #NUM! | 2.7234 | 0.0007 | 2.5813 | 0.0006 | 6.927 | #NUM! | 6.9101 | #NUM! | 7.2671 | #NUM! | 95.322 | #NUM! | 95.09 | #NUM! | 2.7986 | 0.0659 | 3.4109 | 0.087565 | 8.902 | 0.002 | 1.5961 | 0.1901 | 0.0068 | 95.3584 | 0.1614 | 2019-06-29 15:04:24 | 434006 | 8.06 | 0.013 | 8.069 | 8.051 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.06 | 0.018 | FC
FC-20 | 2019-06-29 | 09:34:14 | 39200022 | 434007 | 47 | 0 | -46 | 29 | 30.142 | 0.166 | 6.3255 | 0.0226 | 6.3238 | 0.0211 | 3.3267 | 0.001929 | 3.3266 | 0.001772 | 33.4226 | #NUM! | 33.423 | #NUM! | 26.2652 | #NUM! | 26.2657 | #NUM! | 2.7932 | 0.0006 | 2.654 | 0.0009 | 6.768 | #NUM! | 6.754 | #NUM! | 6.9287 | #NUM! | 97.685 | #NUM! | 97.482 | #NUM! | 1.5083 | 0.0405 | 9.8485 | 0.19197 | 8.903 | 0 | 2.1218 | 0.1829 | 0.0013 | 95.5301 | 0.0309 | 2019-06-29 14:54:13 | 434007 | 7.878 | 0.001 | 7.878 | 7.879 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.8785 | 0.001 | FC
FC-20 | 2019-06-29 | 09:36:01 | 39200022 | 434008 | 47 | 0 | -46 | 29 | 18.798 | 0.171 | 6.9938 | 0.0116 | 6.9929 | 0.0117 | 3.3832 | 0.000962 | 3.3831 | 0.000978 | 33.3993 | #NUM! | 33.4 | #NUM! | 26.1588 | #NUM! | 26.1595 | #NUM! | 2.7926 | 0.0011 | 2.6545 | 0.0004 | 6.6534 | #NUM! | 6.6451 | #NUM! | 6.8235 | #NUM! | 97.511 | #NUM! | 97.389 | #NUM! | 0.5961 | 0.0045 | 28.497 | 0.41785 | 8.894 | 0 | 4.8454 | 0.1302 | 0.0032 | 96.799 | 0.0772 | 2019-06-29 14:42:40 | 434008 | 7.755 | 0.021 | 7.77 | 7.74 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.755 | 0.03 | FC
FC-20 | 2019-06-29 | 09:37:56 | 39200022 | 434009 | 47 | 0 | -46 | 29 | 9.861 | 0.088 | 9.4524 | 0.0094 | 9.4547 | 0.0055 | 3.589 | 0.000782 | 3.5893 | 0.000443 | 33.2556 | #NUM! | 33.2561 | #NUM! | 25.6802 | #NUM! | 25.6802 | #NUM! | 2.8305 | 0.0011 | 2.692 | 0.0006 | 6.3656 | #NUM! | 6.357 | #NUM! | 6.4627 | #NUM! | 98.501 | #NUM! | 98.368 | #NUM! | 0.358 | 0.0094 | 81.803 | 2.711 | 8.86 | 0.003 | 0.3704 | 0.0815 | 0.0002 | 97.9828 | 0.0054 | 2019-06-29 14:31:18 | 434009 | 7.47 | 0.008 | 7.475 | 7.464 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.4695 | 0.011 | FC
FC-20 | 2019-06-29 | 09:39:56 | 39200022 | 434010 | 47 | 0 | -46 | 29 | 4.733 | 0.125 | 9.4995 | 0.0004 | 9.4991 | 0.0009 | 3.5929 | 4e-05 | 3.593 | 7.9e-05 | 33.2545 | #NUM! | 33.2553 | #NUM! | 25.6717 | #NUM! | 25.6724 | #NUM! | 2.8317 | 0.0009 | 2.6952 | 0.0006 | 6.3465 | #NUM! | 6.3494 | #NUM! | 6.4561 | #NUM! | 98.308 | #NUM! | 98.352 | #NUM! | 0.2992 | 0.009 | 166.58 | 18.848 | 8.863 | 0 | 0.4213 | 0.0817 | 0.0005 | 97.9778 | 0.0115 | 2019-06-29 14:19:19 | 434010 | 7.454 | 0.008 | 7.46 | 7.449 | 39200022 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.4545 | 0.011 | FC
FC-24 | 2019-06-29 | 19:11:48 | 39200026 | 434032 | 47 | 0 | -45 | 12.8 | 166.914 | 0.023 | 4.2908 | 0.0007 | 4.2913 | 0.0008 | 3.269 | 8.4e-05 | 3.2691 | 8.9e-05 | 34.7618 | #NUM! | 34.7628 | #NUM! | 27.5659 | #NUM! | 27.5666 | #NUM! | 2.0098 | 0.0004 | 1.9085 | 0.0003 | 4.7457 | #NUM! | 4.7122 | #NUM! | 7.2056 | #NUM! | 65.863 | #NUM! | 65.397 | #NUM! | 0.049 | 0.008 | 0.11769 | 0.000205 | 8.653 | 0.002 | 1.4163 | -0.0339 | 0.0006 | 100.8513 | 0.0159 | 2019-06-30 00:52:58 | 434032 | 5.523 | 0.006 | 5.518 | 5.527 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 5.5225 | 0.009 | FC
FC-24 | 2019-06-29 | 19:13:39 | 39200026 | 434033 | 47 | 0 | -45 | 12.8 | 148.887 | 0.081 | 4.0029 | 0.0004 | 4.0032 | 0.0006 | 3.237 | 4.5e-05 | 3.2371 | 6e-05 | 34.6971 | #NUM! | 34.698 | #NUM! | 27.5449 | #NUM! | 27.5456 | #NUM! | 2.0667 | 0.0006 | 1.9556 | 0.0007 | 4.948 | #NUM! | 4.8994 | #NUM! | 7.259 | #NUM! | 68.165 | #NUM! | 67.495 | #NUM! | 0.0421 | 0.0034 | 0.11771 | 0.00020184 | 8.661 | 0.001 | 1.0446 | -0.0594 | 0.0031 | 101.4963 | 0.0793 | 2019-06-30 00:44:27 | 434033 | 5.804 | 0.041 | 5.775 | 5.833 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 5.804 | 0.058 | FC
FC-24 | 2019-06-29 | 19:16:06 | 39200026 | 434034 | 47 | 0 | -45 | 12.8 | 99.678 | 0.038 | 3.9392 | 0.0064 | 3.9465 | 0.0074 | 3.2085 | 0.000535 | 3.2091 | 0.000645 | 34.4518 | #NUM! | 34.4515 | #NUM! | 27.3563 | #NUM! | 27.3554 | #NUM! | 2.2428 | 0.0017 | 2.1171 | 0.0006 | 5.5115 | #NUM! | 5.4363 | #NUM! | 7.2823 | #NUM! | 75.686 | #NUM! | 74.653 | #NUM! | 0.0304 | 0.0047 | 0.11759 | 0.00019426 | 8.705 | 0.002 | 1.0216 | -0.0986 | 0.0032 | 102.4946 | 0.0827 | 2019-06-30 00:33:47 | 434034 | 6.328 | 0.011 | 6.32 | 6.336 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.328 | 0.016 | FC
FC-24 | 2019-06-29 | 19:18:02 | 39200026 | 434035 | 47 | 0 | -45 | 12.8 | 75.22 | 0.057 | 4.0885 | 0.0075 | 4.0927 | 0.0065 | 3.2135 | 0.000807 | 3.214 | 0.000684 | 34.3676 | #NUM! | 34.3687 | #NUM! | 27.2739 | #NUM! | 27.2743 | #NUM! | 2.2678 | 0.0008 | 2.1461 | 0.0006 | 5.5417 | #NUM! | 5.4921 | #NUM! | 7.2602 | #NUM! | 76.332 | #NUM! | 75.648 | #NUM! | 0.0765 | 0.0037 | 0.40235 | 0.0068653 | 8.713 | 0.001 | 1.1116 | -0.1025 | 0.0011 | 102.5948 | 0.0277 | 2019-06-30 00:25:08 | 434035 | 6.418 | 0.004 | 6.415 | 6.42 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.4175 | 0.005 | FC
FC-24 | 2019-06-29 | 19:19:58 | 39200026 | 434036 | 47 | 0 | -45 | 12.8 | 50.302 | 0.17 | 4.0765 | 0.0012 | 4.076 | 0.0008 | 3.1857 | 5.9e-05 | 3.1857 | 6.8e-05 | 34.064 | #NUM! | 34.0649 | #NUM! | 27.0337 | #NUM! | 27.0344 | #NUM! | 2.5282 | 0.0005 | 2.3845 | 0.0007 | 6.3448 | #NUM! | 6.2886 | #NUM! | 7.2772 | #NUM! | 87.19 | #NUM! | 86.418 | #NUM! | 0.5671 | 0.0152 | 2.0808 | 0.041385 | 8.818 | 0 | 3.1262 | -0.07 | 0.0005 | 101.7649 | 0.0132 | 2019-06-30 00:14:44 | 434036 | 7.354 | 0.006 | 7.358 | 7.35 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.354 | 0.008 | FC
FC-24 | 2019-06-29 | 19:21:24 | 39200026 | 434037 | 47 | 0 | -45 | 12.8 | 39.996 | 0.102 | 4.3174 | 0.0133 | 4.3177 | 0.0128 | 3.1999 | 0.001057 | 3.1999 | 0.000954 | 33.9871 | #NUM! | 33.987 | #NUM! | 26.9474 | #NUM! | 26.9472 | #NUM! | 2.6042 | 0.0018 | 2.4529 | 0.0006 | 6.5535 | #NUM! | 6.4567 | #NUM! | 7.2387 | #NUM! | 90.536 | #NUM! | 89.199 | #NUM! | 1.7059 | 0.1307 | 4.628 | 0.066201 | 8.845 | 0 | 0.6346 | 0.0088 | 0.0027 | 99.7794 | 0.0684 | 2019-06-30 00:03:52 | 434037 | 7.587 | 0.018 | 7.6 | 7.574 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.587 | 0.026 | FC
FC-24 | 2019-06-29 | 19:22:53 | 39200026 | 434038 | 47 | 0 | -45 | 12.8 | 30.878 | 0.091 | 4.9213 | 0.0064 | 4.9191 | 0.0044 | 3.2451 | 0.000471 | 3.2449 | 0.000377 | 33.9003 | #NUM! | 33.9004 | #NUM! | 26.8123 | #NUM! | 26.8126 | #NUM! | 2.7144 | 0.0009 | 2.5585 | 0.0011 | 6.7539 | #NUM! | 6.6765 | #NUM! | 7.1388 | #NUM! | 94.61 | #NUM! | 93.526 | #NUM! | 2.6613 | 0.0105 | 9.1571 | 0.078235 | 8.877 | 0.002 | 1.0227 | 0.156 | 0.0064 | 96.1751 | 0.155 | 2019-06-30 01:01:00 | 434038 | 7.853 | 0.002 | 7.854 | 7.851 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.8525 | 0.003 | FC
FC-24 | 2019-06-29 | 19:24:13 | 39200026 | 434039 | 47 | 0 | -45 | 12.8 | 20.366 | 0.08 | 6.22 | 0.002 | 6.2268 | 0.003 | 3.3454 | 0.000169 | 3.346 | 0.000231 | 33.7402 | #NUM! | 33.7399 | #NUM! | 26.5294 | #NUM! | 26.5282 | #NUM! | 2.8182 | 0.0009 | 2.6517 | 0.0009 | 6.8451 | #NUM! | 6.7436 | #NUM! | 6.9311 | #NUM! | 98.763 | #NUM! | 97.298 | #NUM! | 2.5561 | 0.0238 | 26.831 | 0.36619 | 8.902 | 0.003 | 0.7267 | 0.2792 | 0.0016 | 93.2573 | 0.0385 | 2019-06-30 01:12:39 | 434039 | 7.936 | 0.008 | 7.93 | 7.942 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.936 | 0.012 | FC
FC-24 | 2019-06-29 | 19:25:39 | 39200026 | 434040 | 47 | 0 | -45 | 12.8 | 10.186 | 0.067 | 7.1437 | 0.0175 | 7.1528 | 0.0351 | 3.4162 | 0.00107 | 3.4169 | 0.002054 | 33.6192 | #NUM! | 33.6182 | #NUM! | 26.3115 | #NUM! | 26.3094 | #NUM! | 2.8808 | 0.0006 | 2.713 | 0.0007 | 6.8597 | #NUM! | 6.7737 | #NUM! | 6.7903 | #NUM! | 101.026 | #NUM! | 99.759 | #NUM! | 0.4167 | 0.0374 | 67.956 | 1.2936 | 8.902 | 0 | 1.703 | 0.1748 | 0.0045 | 95.7238 | 0.108 | 2019-06-30 01:24:54 | 434040 | 7.921 | 0.016 | 7.909 | 7.932 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.9205 | 0.023 | FC
FC-24 | 2019-06-29 | 19:26:41 | 39200026 | 434041 | 47 | 0 | -45 | 12.8 | 5.938 | 0.219 | 9.6803 | 0.0394 | 9.7586 | 0.0351 | 3.6273 | 0.003467 | 3.634 | 0.002942 | 33.439 | #NUM! | 33.4354 | #NUM! | 25.7863 | #NUM! | 25.7706 | #NUM! | 2.8845 | 0.0013 | 2.7226 | 0.0006 | 6.4689 | #NUM! | 6.4014 | #NUM! | 6.4229 | #NUM! | 100.721 | #NUM! | 99.67 | #NUM! | 0.2517 | 0.0036 | 134.39 | 5.0377 | 8.864 | 0.002 | -0.1493 | 0.1045 | 0.0005 | 97.4214 | 0.0133 | 2019-06-30 01:37:14 | 434041 | 7.529 | 0.007 | 7.524 | 7.534 | 39200026 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.529 | 0.01 | FC
FC-33 | 2019-06-30 | 10:28:46 | 39200034 | 434075 | 47 | 0 | -43 | 24 | 1205.574 | 0.163 | 3.3756 | 0.0002 | 3.3754 | 0.0001 | 3.2437 | 2.5e-05 | 3.2437 | 1.6e-05 | 34.8866 | #NUM! | 34.8872 | #NUM! | 27.759 | #NUM! | 27.7595 | #NUM! | 2.1301 | 0.0005 | 2.0239 | 0.0005 | 6.0121 | #NUM! | 5.9851 | #NUM! | 7.3612 | #NUM! | 81.674 | #NUM! | 81.308 | #NUM! | 0.0117 | 0.0055 | 0.11776 | 0.00018289 | 8.727 | 0 | 1.8385 | -0.1119 | 0.0002 | 102.8383 | 0.0066 | 2019-06-30 15:10:42 | 434075 | 7.053 | 0.001 | 7.054 | 7.053 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.0535 | 0.001 | FC
FC-33 | 2019-06-30 | 10:36:31 | 39200034 | 434076 | 47 | 0 | -43 | 24 | 1000.237 | 0.318 | 3.3956 | 0.0001 | 3.3953 | 0.0001 | 3.2359 | 2.5e-05 | 3.2359 | 2.5e-05 | 34.8772 | #NUM! | 34.8778 | #NUM! | 27.7496 | #NUM! | 27.7501 | #NUM! | 2.1983 | 0.0006 | 2.0881 | 0.0004 | 6.0943 | #NUM! | 6.0718 | #NUM! | 7.3581 | #NUM! | 82.826 | #NUM! | 82.519 | #NUM! | 0.0039 | 0.0039 | 0.11769 | 0.000205 | 8.741 | 0 | 2.9052 | -0.1107 | 0.0005 | 102.8052 | 0.0124 | 2019-06-30 15:22:44 | 434076 | 7.131 | 0.001 | 7.132 | 7.13 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.131 | 0.002 | FC
FC-33 | 2019-06-30 | 10:51:26 | 39200034 | 434077 | 47 | 0 | -43 | 24 | 500.315 | 0.062 | 3.8132 | 0.0046 | 3.8185 | 0.0093 | 3.2515 | 0.000455 | 3.2521 | 0.00083 | 34.8824 | #NUM! | 34.8834 | #NUM! | 27.712 | #NUM! | 27.7122 | #NUM! | 2.2786 | 0.0005 | 2.1611 | 0.0008 | 5.9058 | #NUM! | 5.8664 | #NUM! | 7.2834 | #NUM! | 81.089 | #NUM! | 80.547 | #NUM! | 0.0186 | 0.0036 | 0.1177 | 0.0002036 | 8.762 | 0.001 | 2.1224 | -0.1132 | 0.0004 | 102.8711 | 0.0096 | 2019-06-30 15:33:49 | 434077 | 6.87 | 0.005 | 6.874 | 6.867 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.8705 | 0.007 | FC
FC-33 | 2019-06-30 | 10:58:53 | 39200034 | 434078 | 47 | 0 | -43 | 24 | 250.117 | 0.072 | 4.8047 | 0.0007 | 4.8028 | 0.001 | 3.3306 | 6.3e-05 | 3.3305 | 8.5e-05 | 34.8989 | #NUM! | 34.8996 | #NUM! | 27.6177 | #NUM! | 27.6185 | #NUM! | 2.1719 | 0.0005 | 2.0627 | 0.0006 | 5.2359 | #NUM! | 5.211 | #NUM! | 7.1109 | #NUM! | 73.633 | #NUM! | 73.284 | #NUM! | 0.0197 | 0.0033 | 0.11769 | 0.000205 | 8.751 | 0 | 0.8588 | -0.1182 | 0.0003 | 103 | 0.0083 | 2019-06-30 15:44:16 | 434078 | 6.122 | 0.006 | 6.117 | 6.126 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.1215 | 0.009 | FC
FC-33 | 2019-06-30 | 11:02:38 | 39200034 | 434079 | 47 | 0 | -43 | 24 | 149.663 | 0.186 | 5.1104 | 0.0016 | 5.1106 | 0.0019 | 3.3409 | 0.000145 | 3.3409 | 0.000171 | 34.751 | #NUM! | 34.7516 | #NUM! | 27.4649 | #NUM! | 27.4654 | #NUM! | 2.1957 | 0.0006 | 2.0875 | 0.0007 | 5.2057 | #NUM! | 5.1903 | #NUM! | 7.0665 | #NUM! | 73.67 | #NUM! | 73.452 | #NUM! | 0.0244 | 0.0049 | 0.11767 | 0.00020708 | 8.75 | 0.002 | 2.123 | -0.1126 | 0.0008 | 102.8539 | 0.0201 | 2019-06-30 15:53:23 | 434079 | 6.08 | 0.008 | 6.074 | 6.086 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.08 | 0.012 | FC
FC-33 | 2019-06-30 | 11:05:51 | 39200034 | 434080 | 47 | 0 | -43 | 24 | 100.669 | 0.168 | 5.802 | 0.0035 | 5.8006 | 0.0057 | 3.4001 | 0.000365 | 3.4001 | 0.00067 | 34.7387 | #NUM! | 34.7396 | #NUM! | 27.3713 | #NUM! | 27.3722 | #NUM! | 2.2031 | 0.0007 | 2.096 | 0.0011 | 5.1016 | #NUM! | 5.0681 | #NUM! | 6.9528 | #NUM! | 73.377 | #NUM! | 72.894 | #NUM! | 0.0293 | 0.0062 | 0.11769 | 0.00020604 | 8.754 | 0.001 | 1.2277 | -0.1101 | 0.0005 | 102.7913 | 0.0123 | 2019-06-30 16:04:17 | 434080 | 5.985 | 0.022 | 6 | 5.969 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 5.9845 | 0.031 | FC
FC-33 | 2019-06-30 | 11:07:54 | 39200034 | 434081 | 47 | 0 | -43 | 24 | 74.738 | 0.176 | 5.7798 | 0.0141 | 5.785 | 0.0036 | 3.3865 | 0.001759 | 3.3871 | 0.000444 | 34.6201 | #NUM! | 34.6225 | #NUM! | 27.2802 | #NUM! | 27.2815 | #NUM! | 2.3497 | 0.0011 | 2.2257 | 0.0014 | 5.5156 | #NUM! | 5.4668 | #NUM! | 6.962 | #NUM! | 79.228 | #NUM! | 78.527 | #NUM! | 0.0522 | 0.0053 | 0.11768 | 0.00020674 | 8.788 | 0 | 0.9547 | -0.1014 | 0.0013 | 102.5676 | 0.0328 | 2019-06-30 17:04:47 | 434081 | 6.476 | 0.009 | 6.47 | 6.483 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.4765 | 0.013 | FC
FC-33 | 2019-06-30 | 11:09:56 | 39200034 | 434082 | 47 | 0 | -43 | 24 | 49.942 | 0.291 | 6.5192 | 0.0081 | 6.5238 | 0.0064 | 3.4523 | 0.000897 | 3.4528 | 0.000726 | 34.6188 | #NUM! | 34.6202 | #NUM! | 27.1838 | #NUM! | 27.1843 | #NUM! | 2.4862 | 0.0009 | 2.3524 | 0.0007 | 5.7951 | #NUM! | 5.7637 | #NUM! | 6.8433 | #NUM! | 84.686 | #NUM! | 84.228 | #NUM! | 0.2062 | 0.0183 | 0.16665 | 0.0071047 | 8.827 | 0.002 | 1.1954 | -0.0804 | 0.0015 | 102.0306 | 0.0369 | 2019-06-30 16:54:06 | 434082 | 6.778 | 0.011 | 6.786 | 6.771 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.7785 | 0.015 | FC
FC-33 | 2019-06-30 | 11:11:55 | 39200034 | 434083 | 47 | 0 | -43 | 24 | 39.177 | 0.076 | 6.4549 | 0.0014 | 6.4554 | 0.0028 | 3.4328 | 0.00014 | 3.4329 | 0.000246 | 34.4721 | #NUM! | 34.4725 | #NUM! | 27.0766 | #NUM! | 27.0769 | #NUM! | 2.5675 | 0.0009 | 2.4323 | 0.0005 | 6.0432 | #NUM! | 6.009 | #NUM! | 6.8601 | #NUM! | 88.095 | #NUM! | 87.596 | #NUM! | 0.9452 | 0.0332 | 0.85743 | 0.014955 | 8.851 | 0.002 | 3.1092 | 0.0031 | 0.0023 | 99.9234 | 0.0579 | 2019-06-30 16:43:56 | 434083 | 7.048 | 0.013 | 7.038 | 7.057 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.0475 | 0.019 | FC
FC-33 | 2019-06-30 | 11:14:18 | 39200034 | 434084 | 47 | 0 | -43 | 24 | 19.932 | 0.335 | 8.6753 | 0.0322 | 8.6826 | 0.0597 | 3.6215 | 0.001747 | 3.622 | 0.003345 | 34.3277 | #NUM! | 34.3265 | #NUM! | 26.6429 | #NUM! | 26.6408 | #NUM! | 3.0689 | 0.0015 | 2.8968 | 0.003 | 7.1108 | #NUM! | 7.0727 | #NUM! | 6.5293 | #NUM! | 108.911 | #NUM! | 108.327 | #NUM! | 7.2053 | 0.1655 | 23.811 | 1.7258 | 8.967 | 0.001 | 1.0622 | 0.8999 | 0.0194 | 79.8553 | 0.3883 | 2019-06-30 16:32:43 | 434084 | 8.248 | 0.032 | 8.226 | 8.271 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.2485 | 0.045 | FC
FC-33 | 2019-06-30 | 11:17:54 | 39200034 | 434086 | 47 | 0 | -43 | 24 | 5.162 | 0.043 | 9.5694 | 0.0035 | 9.5688 | 0.0024 | 3.6718 | 0.000303 | 3.6718 | 0.000211 | 34.0014 | #NUM! | 34.0021 | #NUM! | 26.2441 | #NUM! | 26.2447 | #NUM! | 2.9749 | 0.0007 | 2.819 | 0.0011 | 6.6978 | #NUM! | 6.6781 | #NUM! | 6.4153 | #NUM! | 104.409 | #NUM! | 104.102 | #NUM! | 0.717 | 0.0029 | 418.97 | 26.337 | 8.947 | 0 | 3.3049 | 0.531 | 0.0028 | 87.569 | 0.0622 | 2019-06-30 16:15:12 | 434086 | 7.832 | 0.034 | 7.856 | 7.808 | 39200034 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.832 | 0.048 | FC
BB-14 | 2019-07-01 | 23:56:09 | 39200045 | 434099 | 50 | 10.6 | -48 | 28.3 | 1206.319 | 0.131 | 3.3506 | 0.0001 | 3.3504 | 0.0001 | 3.2406 | 1.8e-05 | 3.2406 | 1.4e-05 | 34.8751 | #NUM! | 34.8762 | #NUM! | 27.7523 | #NUM! | 27.7531 | #NUM! | 2.1516 | 0.0003 | 2.0497 | 0.0007 | 6.0932 | #NUM! | 6.0804 | #NUM! | 7.3663 | #NUM! | 82.719 | #NUM! | 82.544 | #NUM! | 0.0131 | 0.0035 | 0.11769 | 0.000205 | 8.718 | 0.002 | 2.7901 | -0.1078 | 0.0003 | 102.7321 | 0.0068 | 2019-07-02 06:30:33 | 434099 | 7.23 | 0.014 | 7.24 | 7.22 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.23 | 0.02 | BB
BB-14 | 2019-07-02 | 00:02:19 | 39200045 | 434100 | 50 | 10.6 | -48 | 28.3 | 1000.009 | 0.129 | 3.3696 | 0.0002 | 3.3694 | 0.0002 | 3.2329 | 2.5e-05 | 3.233 | 2e-05 | 34.8693 | #NUM! | 34.8705 | #NUM! | 27.7458 | #NUM! | 27.7467 | #NUM! | 2.2102 | 0.0006 | 2.1024 | 0.0006 | 6.1439 | #NUM! | 6.1252 | #NUM! | 7.3632 | #NUM! | 83.442 | #NUM! | 83.188 | #NUM! | 0.014 | 0.004 | 0.11774 | 0.00019426 | 8.728 | 0.002 | 1.9645 | -0.1067 | 0.0002 | 102.705 | 0.0039 | 2019-07-02 06:40:45 | 434100 | 7.285 | 0.052 | 7.248 | 7.322 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.285 | 0.074 | BB
BB-14 | 2019-07-02 | 00:13:45 | 39200045 | 434101 | 50 | 10.6 | -48 | 28.3 | 498.38 | 0.179 | 3.37 | 0.0001 | 3.3696 | 0.0001 | 3.2091 | 2.6e-05 | 3.2092 | 1.6e-05 | 34.8476 | #NUM! | 34.8489 | #NUM! | 27.7285 | #NUM! | 27.7295 | #NUM! | 2.344 | 0.0005 | 2.2305 | 0.0008 | 6.2049 | #NUM! | 6.1931 | #NUM! | 7.3642 | #NUM! | 84.26 | #NUM! | 84.099 | #NUM! | 0.0189 | 0.0041 | 0.11774 | 0.00019092 | 8.752 | 0.002 | 1.0512 | -0.1057 | 0.0002 | 102.6762 | 0.0054 | 2019-07-02 06:51:03 | 434101 | 7.299 | 0.031 | 7.321 | 7.277 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.299 | 0.044 | BB
BB-14 | 2019-07-02 | 00:20:49 | 39200045 | 434102 | 50 | 10.6 | -48 | 28.3 | 249.777 | 0.264 | 3.3376 | 0.0001 | 3.3373 | 0.0001 | 3.1923 | 2.1e-05 | 3.1924 | 2.5e-05 | 34.8149 | #NUM! | 34.8162 | #NUM! | 27.7055 | #NUM! | 27.7066 | #NUM! | 2.4222 | 0.0007 | 2.3023 | 0.0006 | 6.2713 | #NUM! | 6.2575 | #NUM! | 7.3717 | #NUM! | 85.075 | #NUM! | 84.887 | #NUM! | 0.053 | 0.029 | 0.11775 | 0.00018714 | 8.761 | 0.001 | 0.8155 | -0.1037 | 0.0003 | 102.6258 | 0.0083 | 2019-07-02 07:01:55 | 434102 | 7.405 | 0.039 | 7.432 | 7.377 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.4045 | 0.055 | BB
BB-14 | 2019-07-02 | 00:24:36 | 39200045 | 434103 | 50 | 10.6 | -48 | 28.3 | 150.049 | 0.152 | 3.2956 | 0.0008 | 3.2955 | 0.0015 | 3.1799 | 6.2e-05 | 3.18 | 0.000119 | 34.7654 | #NUM! | 34.7667 | #NUM! | 27.67 | #NUM! | 27.6711 | #NUM! | 2.4834 | 0.0006 | 2.3631 | 0.0002 | 6.3916 | #NUM! | 6.3989 | #NUM! | 7.3818 | #NUM! | 86.589 | #NUM! | 86.687 | #NUM! | 0.1448 | 0.0293 | 0.11774 | 0.00019092 | 8.77 | 0.001 | 0.7004 | -0.1035 | 0.0003 | 102.6208 | 0.0075 | 2019-07-02 07:14:37 | 434103 | 7.488 | 0.015 | 7.498 | 7.477 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.4875 | 0.021 | BB
BB-14 | 2019-07-02 | 00:27:29 | 39200045 | 434104 | 50 | 10.6 | -48 | 28.3 | 99.688 | 0.129 | 3.3993 | 0.0016 | 3.3988 | 0.0023 | 3.1843 | 0.000126 | 3.1843 | 0.000199 | 34.7344 | #NUM! | 34.7355 | #NUM! | 27.6353 | #NUM! | 27.6362 | #NUM! | 2.4653 | 0.0008 | 2.3453 | 0.0003 | 6.2831 | #NUM! | 6.2792 | #NUM! | 7.3646 | #NUM! | 85.316 | #NUM! | 85.263 | #NUM! | 0.2382 | 0.0394 | 0.11768 | 0.00020674 | 8.767 | 0.002 | 1.0178 | -0.0973 | 0.0022 | 102.4624 | 0.0567 | 2019-07-02 07:27:40 | 434104 | 7.32 | 0.008 | 7.314 | 7.326 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.32 | 0.012 | BB
BB-14 | 2019-07-02 | 00:29:17 | 39200045 | 434105 | 50 | 10.6 | -48 | 28.3 | 74.474 | 0.157 | 3.4942 | 0.0007 | 3.4944 | 0.0003 | 3.1872 | 5.2e-05 | 3.1873 | 3e-05 | 34.6817 | #NUM! | 34.6828 | #NUM! | 27.5841 | #NUM! | 27.585 | #NUM! | 2.5102 | 0.0004 | 2.3819 | 0.0004 | 6.3825 | #NUM! | 6.3669 | #NUM! | 7.3501 | #NUM! | 86.836 | #NUM! | 86.624 | #NUM! | 0.6962 | 0.0619 | 0.1176 | 0.00019718 | 8.773 | 0.002 | 2.4946 | -0.0513 | 0.0279 | 101.2931 | 0.707 | 2019-07-02 05:30:57 | 434105 | 7.523 | 0.035 | 7.548 | 7.498 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.523 | 0.05 | BB
BB-14 | 2019-07-02 | 00:31:02 | 39200045 | 434106 | 50 | 10.6 | -48 | 28.3 | 50.729 | 0.087 | 4.139 | 0.0028 | 4.1292 | 0.0089 | 3.2409 | 0.000181 | 3.2402 | 0.000699 | 34.6524 | #NUM! | 34.6546 | #NUM! | 27.4951 | #NUM! | 27.4979 | #NUM! | 2.5551 | 0.0014 | 2.4254 | 0.0004 | 6.4064 | #NUM! | 6.3777 | #NUM! | 7.2374 | #NUM! | 88.52 | #NUM! | 88.124 | #NUM! | 3.5219 | 0.4754 | 0.11751 | 0.00013579 | 8.783 | 0.002 | 1.7157 | 0.1142 | 0.0017 | 97.1861 | 0.0407 | 2019-07-02 05:42:43 | 434106 | 7.622 | 0.077 | 7.677 | 7.568 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.6225 | 0.109 | BB
BB-14 | 2019-07-02 | 00:32:26 | 39200045 | 434107 | 50 | 10.6 | -48 | 28.3 | 40.502 | 0.018 | 5.0369 | 0.0025 | 5.0366 | 0.0031 | 3.3154 | 0.000161 | 3.3154 | 0.000193 | 34.5898 | #NUM! | 34.5909 | #NUM! | 27.3458 | #NUM! | 27.3467 | #NUM! | 2.6703 | 0.0014 | 2.5338 | 0.0003 | 6.5727 | #NUM! | 6.5705 | #NUM! | 7.0864 | #NUM! | 92.752 | #NUM! | 92.722 | #NUM! | 10.1539 | 0.5324 | 0.11749 | 0.00011342 | 8.84 | 0.001 | 1.1313 | 0.4611 | 0.0105 | 89.1126 | 0.2348 | 2019-07-02 05:56:25 | 434107 | 7.724 | 0.022 | 7.74 | 7.709 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.7245 | 0.031 | BB
BB-14 | 2019-07-02 | 00:33:58 | 39200045 | 434108 | 50 | 10.6 | -48 | 28.3 | 20.037 | 0.046 | 6.0501 | 0.0018 | 6.0514 | 0.0016 | 3.3961 | 0.000202 | 3.3963 | 0.000237 | 34.4811 | #NUM! | 34.4817 | #NUM! | 27.1362 | #NUM! | 27.1366 | #NUM! | 3.0448 | 0.0007 | 2.8899 | 0.0005 | 7.508 | #NUM! | 7.4966 | #NUM! | 6.9245 | #NUM! | 108.429 | #NUM! | 108.265 | #NUM! | 9.8662 | 0.7458 | 0.11762 | 0.00020184 | 8.98 | 0.002 | 2.1285 | 0.6812 | 0.0118 | 84.3411 | 0.2486 | 2019-07-02 06:08:09 | 434108 | 8.791 | 0.001 | 8.792 | 8.79 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.791 | 0.002 | BB
BB-14 | 2019-07-02 | 00:37:02 | 39200045 | 434110 | 50 | 10.6 | -48 | 28.3 | 5.729 | 0.06 | 9.4671 | 0.0521 | 9.5116 | 0.0607 | 3.7061 | 0.004151 | 3.7101 | 0.004487 | 34.4531 | #NUM! | 34.4516 | #NUM! | 26.6141 | #NUM! | 26.6056 | #NUM! | 3.2394 | 0.0009 | 3.0768 | 0.0019 | 7.4139 | #NUM! | 7.3872 | #NUM! | 6.411 | #NUM! | 115.649 | #NUM! | 115.233 | #NUM! | 2.7241 | 0.156 | 0.11751 | 0.00013579 | 9.008 | 0.001 | 34.7795 | 0.5236 | 0.0012 | 87.7306 | 0.0262 | 2019-07-02 06:19:14 | 434110 | 8.675 | 0.024 | 8.658 | 8.692 | 39200045 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.675 | 0.034 | BB
BB-10 | 2019-07-02 | 17:39:51 | 39200049 | 434147 | 49 | 31 | -50 | 32 | 332.258 | 0.106 | 3.1994 | 0.0001 | 3.1992 | 0.0002 | 3.1759 | 1.8e-05 | 3.176 | 2.1e-05 | 34.7197 | #NUM! | 34.7209 | #NUM! | 27.6428 | #NUM! | 27.6438 | #NUM! | 2.3117 | 0.0008 | 2.198 | 0.0006 | 6.0042 | #NUM! | 5.9834 | #NUM! | 7.4015 | #NUM! | 81.123 | #NUM! | 80.842 | #NUM! | 0.0344 | 0.0041 | 0.11766 | 0.00020708 | 8.728 | 0 | 1.8292 | 0.0835 | 0.0015 | 97.9335 | 0.038 | 2019-07-02 23:34:51 | 434147 | 7.043 | 0.002 | 7.045 | 7.042 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.0435 | 0.003 | BB
BB-10 | 2019-07-02 | 17:45:41 | 39200049 | 434148 | 49 | 31 | -50 | 32 | 149.687 | 0.075 | 1.9192 | 0.0001 | 1.919 | 0.0001 | 3.0116 | 2.3e-05 | 3.0117 | 1.8e-05 | 34.1714 | #NUM! | 34.1728 | #NUM! | 27.314 | #NUM! | 27.3152 | #NUM! | 2.4141 | 0.0007 | 2.2904 | 0.0004 | 6.4486 | #NUM! | 6.4302 | #NUM! | 7.6697 | #NUM! | 84.08 | #NUM! | 83.84 | #NUM! | 0.0232 | 0.0052 | 0.11767 | 0.00020708 | 8.761 | 0.002 | 2.0802 | -0.0638 | 0.0007 | 101.6076 | 0.0175 | 2019-07-02 23:44:19 | 434148 | 7.531 | 0.003 | 7.529 | 7.533 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.531 | 0.004 | BB
BB-10 | 2019-07-02 | 17:48:49 | 39200049 | 434149 | 49 | 31 | -50 | 32 | 98.942 | 0.362 | 0.7961 | 0.0051 | 0.7954 | 0.0054 | 2.8826 | 0.000568 | 2.8827 | 0.000615 | 33.7754 | #NUM! | 33.7767 | #NUM! | 27.0734 | #NUM! | 27.0745 | #NUM! | 2.4203 | 0.0008 | 2.2942 | 0.0005 | 6.6472 | #NUM! | 6.6422 | #NUM! | 7.914 | #NUM! | 83.994 | #NUM! | 83.93 | #NUM! | 0.0572 | 0.006 | 0.20445 | 0.0093113 | 8.767 | 0 | 0.713 | -0.0617 | 0.0006 | 101.5539 | 0.0146 | 2019-07-02 23:54:31 | 434149 | 7.778 | 0.004 | 7.775 | 7.781 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.778 | 0.006 | BB
BB-10 | 2019-07-02 | 17:52:03 | 39200049 | 434150 | 49 | 31 | -50 | 32 | 76.236 | 0.176 | 0.1927 | 0.0008 | 0.1934 | 0.0015 | 2.8058 | 0.000149 | 2.806 | 0.000234 | 33.4446 | #NUM! | 33.4465 | #NUM! | 26.8409 | #NUM! | 26.8424 | #NUM! | 2.4475 | 0.0006 | 2.3164 | 0.0005 | 6.8657 | #NUM! | 6.8395 | #NUM! | 8.0573 | #NUM! | 85.211 | #NUM! | 84.886 | #NUM! | 0.1062 | 0.0073 | 0.92594 | 0.014937 | 8.777 | 0.002 | 3.0001 | -0.0483 | 0.0007 | 101.2156 | 0.0163 | 2019-07-03 00:04:16 | 434150 | 7.995 | 0.03 | 8.016 | 7.974 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.995 | 0.042 | BB
BB-10 | 2019-07-02 | 17:54:07 | 39200049 | 434151 | 49 | 31 | -50 | 32 | 48.442 | 0.047 | 1.4787 | 0.0011 | 1.4762 | 0.002 | 2.8784 | 8.3e-05 | 2.8783 | 0.000147 | 33.0199 | #NUM! | 33.021 | #NUM! | 26.4225 | #NUM! | 26.4235 | #NUM! | 2.7056 | 0.0007 | 2.5561 | 0.0007 | 7.4748 | #NUM! | 7.4302 | #NUM! | 7.8177 | #NUM! | 95.616 | #NUM! | 95.044 | #NUM! | 0.1575 | 0.0052 | 5.5469 | 0.02077 | 8.903 | 0 | 1.8604 | -0.0306 | 0.004 | 100.7671 | 0.1 | 2019-07-03 00:14:29 | 434151 | 8.748 | 0.028 | 8.729 | 8.768 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.7485 | 0.039 | BB
BB-10 | 2019-07-02 | 17:56:11 | 39200049 | 434152 | 49 | 31 | -50 | 32 | 40.3 | 0.384 | 3.3515 | 0.003 | 3.354 | 0.0044 | 3.0257 | 0.000346 | 3.026 | 0.000496 | 32.9043 | #NUM! | 32.9059 | #NUM! | 26.181 | #NUM! | 26.1821 | #NUM! | 2.7616 | 0.0006 | 2.6169 | 0.0006 | 7.2664 | #NUM! | 7.2556 | #NUM! | 7.4655 | #NUM! | 97.335 | #NUM! | 97.191 | #NUM! | 0.3865 | 0.0174 | 10.851 | 0.36814 | 8.938 | 0.001 | 1.6521 | -0.0019 | 0.0007 | 100.049 | 0.0172 | 2019-07-03 00:24:51 | 434152 | 8.535 | 0.015 | 8.525 | 8.546 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.5355 | 0.021 | BB
BB-10 | 2019-07-02 | 17:58:04 | 39200049 | 434153 | 49 | 31 | -50 | 32 | 30.995 | 0.148 | 3.1521 | 0.0053 | 3.1512 | 0.0064 | 3.0043 | 0.000454 | 3.0043 | 0.000543 | 32.8545 | #NUM! | 32.8557 | #NUM! | 26.1593 | #NUM! | 26.1603 | #NUM! | 2.7525 | 0.0006 | 2.6068 | 0.0008 | 7.2687 | #NUM! | 7.2354 | #NUM! | 7.5048 | #NUM! | 96.856 | #NUM! | 96.412 | #NUM! | 0.683 | 0.0112 | 25.784 | 0.28382 | 8.934 | 0.001 | 1.1389 | 0.0294 | 0.0009 | 99.2684 | 0.0218 | 2019-07-02 23:24:29 | 434153 | 8.541 | 0.007 | 8.536 | 8.546 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.541 | 0.01 | BB
BB-10 | 2019-07-02 | 18:00:05 | 39200049 | 434154 | 49 | 31 | -50 | 32 | 20.129 | 0.31 | 3.8947 | 0.1953 | 3.8353 | 0.1826 | 3.0659 | 0.015146 | 3.0625 | 0.014932 | 32.8465 | #NUM! | 32.8654 | #NUM! | 26.0837 | #NUM! | 26.1045 | #NUM! | 2.7937 | 0.004 | 2.6424 | 0.0047 | 7.2077 | #NUM! | 7.1354 | #NUM! | 7.3697 | #NUM! | 97.804 | #NUM! | 96.823 | #NUM! | 0.6456 | 0.0488 | 61.154 | 2.9805 | 8.933 | 0.003 | 1.4448 | 0.0638 | 0.0046 | 98.4168 | 0.1123 | 2019-07-02 23:13:24 | 434154 | 8.442 | 0.004 | 8.444 | 8.439 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.4415 | 0.005 | BB
BB-10 | 2019-07-02 | 18:02:15 | 39200049 | 434155 | 49 | 31 | -50 | 32 | 10.083 | 0.236 | 4.8017 | 0.001 | 4.8017 | 0.0006 | 3.1162 | 0.000113 | 3.1162 | 6.9e-05 | 32.5389 | #NUM! | 32.5398 | #NUM! | 25.7461 | #NUM! | 25.7468 | #NUM! | 2.8201 | 0.0007 | 2.6698 | 0.0006 | 7.1509 | #NUM! | 7.1279 | #NUM! | 7.2249 | #NUM! | 98.979 | #NUM! | 98.66 | #NUM! | 0.2366 | 0.0039 | 94.973 | 2.4509 | 8.938 | 0.002 | 0.4702 | 0.0441 | 0.0009 | 98.9029 | 0.0226 | 2019-07-02 23:02:27 | 434155 | 8.429 | 0.02 | 8.443 | 8.415 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.429 | 0.028 | BB
BB-10 | 2019-07-02 | 18:04:07 | 39200049 | 434156 | 49 | 31 | -50 | 32 | 5.108 | 0.103 | 4.8565 | 0.0064 | 4.8496 | 0.0055 | 3.1193 | 0.000387 | 3.119 | 0.000354 | 32.5246 | #NUM! | 32.5275 | #NUM! | 25.7288 | #NUM! | 25.7319 | #NUM! | 2.8219 | 0.0011 | 2.6754 | 0.0006 | 7.1336 | #NUM! | 7.1265 | #NUM! | 7.2161 | #NUM! | 98.859 | #NUM! | 98.761 | #NUM! | 0.1918 | 0.0128 | 198.07 | 12.146 | 8.942 | 0.001 | -0.0194 | 0.0413 | 0.0016 | 98.9722 | 0.0392 | 2019-07-02 22:51:39 | 434156 | 8.424 | 0.034 | 8.448 | 8.4 | 39200049 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.424 | 0.048 | BB
BB-06 | 2019-07-03 | 04:27:57 | 39200053 | 434191 | 49 | 6 | -51 | 49.8 | 286.47 | 0.283 | 3.0012 | 0.0003 | 3.0016 | 0.0005 | 3.1442 | 1.8e-05 | 3.1443 | 2.5e-05 | 34.5717 | #NUM! | 34.5722 | #NUM! | 27.5431 | #NUM! | 27.5434 | #NUM! | 2.1844 | 0.0005 | 2.0865 | 0.0005 | 5.5799 | #NUM! | 5.5967 | #NUM! | 7.4452 | #NUM! | 74.948 | #NUM! | 75.173 | #NUM! | 0.0628 | 0.0044 | 0.11777 | 0.00017814 | 8.682 | 0.002 | 0.4931 | 0.0551 | 0.0016 | 98.6322 | 0.0388 | 2019-07-03 10:53:18 | 434191 | 6.651 | 0.003 | 6.653 | 6.649 | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 6.651 | 0.004 | BB
BB-06 | 2019-07-03 | 04:31:58 | 39200053 | 434192 | 49 | 6 | -51 | 49.8 | 150.085 | 0.075 | 0.223 | 0.004 | 0.2277 | 0.0061 | 2.8283 | 0.000407 | 2.8288 | 0.00061 | 33.6628 | #NUM! | 33.6648 | #NUM! | 27.0153 | #NUM! | 27.0166 | #NUM! | 2.3852 | 0.0005 | 2.2633 | 0.0004 | 6.6959 | #NUM! | 6.6849 | #NUM! | 8.0386 | #NUM! | 83.296 | #NUM! | 83.16 | #NUM! | 0.0231 | 0.0058 | 0.11769 | 0.000205 | 8.759 | 0.002 | 0.4071 | -0.0811 | 0.0004 | 102.0489 | 0.0101 | 2019-07-03 10:43:17 | 434192 | 7.931 | 0.032 | 7.909 | 7.954 | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.9315 | 0.045 | BB
BB-06 | 2019-07-03 | 04:34:22 | 39200053 | 434193 | 49 | 6 | -51 | 49.8 | 99.965 | 0.542 | -0.9392 | 0.0053 | -0.9407 | 0.0062 | 2.694 | 0.000598 | 2.6939 | 0.00071 | 33.1785 | #NUM! | 33.1786 | #NUM! | 26.675 | #NUM! | 26.6751 | #NUM! | 2.4145 | 0.0006 | 2.2832 | 0.0006 | 7.022 | #NUM! | 6.9885 | #NUM! | 8.3169 | #NUM! | 84.43 | #NUM! | 84.028 | #NUM! | 0.0373 | 0.0082 | 0.11773 | 0.00019718 | 8.774 | 0 | 1.0041 | -0.0621 | 0.0003 | 101.5656 | 0.0068 | 2019-07-03 10:28:34 | 434193 | 8.231 | 0.103 | 8.304 | 8.158 | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.231 | 0.146 | BB
BB-06 | 2019-07-03 | 04:36:07 | 39200053 | 434194 | 49 | 6 | -51 | 49.8 | 75.77 | 0.37 | -1.0507 | 0.0053 | -1.0507 | 0.0053 | 2.6719 | 0.000777 | 2.6719 | 0.000759 | 33.016 | #NUM! | 33.0164 | #NUM! | 26.5471 | #NUM! | 26.5474 | #NUM! | 2.4337 | 0.0006 | 2.2965 | 0.0004 | 7.1083 | #NUM! | 7.0573 | #NUM! | 8.3513 | #NUM! | 85.116 | #NUM! | 84.506 | #NUM! | 0.0229 | 0.0051 | 0.11778 | 0.00017284 | 8.774 | 0.001 | 1.2453 | -0.0503 | 0.0002 | 101.2654 | 0.0054 | 2019-07-03 10:12:56 | 434194 | 8.344 | 0.033 | 8.32 | 8.367 | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.3435 | 0.047 | BB
BB-06 | 2019-07-03 | 04:38:01 | 39200053 | 434195 | 49 | 6 | -51 | 49.8 | 50.898 | 0.198 | -0.6133 | 0.0026 | -0.6137 | 0.0021 | 2.6926 | 0.000194 | 2.6926 | 0.000151 | 32.8321 | #NUM! | 32.8327 | #NUM! | 26.3825 | #NUM! | 26.3831 | #NUM! | 2.5115 | 0.0004 | 2.3671 | 0.0002 | 7.2803 | #NUM! | 7.2196 | #NUM! | 8.2653 | #NUM! | 88.083 | #NUM! | 87.348 | #NUM! | 0.0253 | 0.007 | 0.1177 | 0.0002036 | 8.781 | 0 | 0.77 | -0.0428 | 0.0002 | 101.0772 | 0.0039 | 2019-07-03 10:06:54 | 434195 | 8.548 | 0 | 8.548 | #NUM! | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.548 | #NUM! | BB
BB-06 | 2019-07-03 | 04:39:36 | 39200053 | 434196 | 49 | 6 | -51 | 49.8 | 40.275 | 0.177 | -0.3236 | 0.0536 | -0.3041 | 0.051 | 2.7093 | 0.004625 | 2.7109 | 0.00432 | 32.7492 | #NUM! | 32.7497 | #NUM! | 26.3038 | #NUM! | 26.3034 | #NUM! | 2.5699 | 0.0006 | 2.4217 | 0.0006 | 7.4139 | #NUM! | 7.3458 | #NUM! | 8.207 | #NUM! | 90.337 | #NUM! | 89.506 | #NUM! | 0.0427 | 0.0044 | 0.11774 | 0.00019426 | 8.784 | 0.001 | 2.7463 | -0.0408 | 0.0003 | 101.0252 | 0.0083 | 2019-07-03 09:42:23 | 434196 | 8.687 | 0.016 | 8.698 | 8.676 | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.687 | 0.022 | BB
BB-06 | 2019-07-03 | 04:41:14 | 39200053 | 434197 | 49 | 6 | -51 | 49.8 | 30.935 | 0.087 | 0.4542 | 0.0012 | 0.4548 | 0.0018 | 2.7674 | 8.2e-05 | 2.7675 | 0.000134 | 32.6874 | #NUM! | 32.6875 | #NUM! | 26.2169 | #NUM! | 26.2169 | #NUM! | 2.6788 | 0.0007 | 2.5249 | 0.0005 | 7.6229 | #NUM! | 7.5564 | #NUM! | 8.0451 | #NUM! | 94.752 | #NUM! | 93.926 | #NUM! | 0.0681 | 0.007 | 0.11776 | 0.00018289 | 8.827 | 0 | 0.6379 | -0.0428 | 0 | 101.0766 | 0 | 2019-07-03 11:03:31 | 434197 | 8.311 | 0.675 | 7.834 | 8.789 | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.3115 | 0.955 | BB
BB-06 | 2019-07-03 | 04:42:47 | 39200053 | 434198 | 49 | 6 | -51 | 49.8 | 20.568 | 0.062 | 3.5922 | 0.0127 | 3.5871 | 0.013 | 3.0048 | 0.000884 | 3.0045 | 0.000884 | 32.4226 | #NUM! | 32.4236 | #NUM! | 25.7752 | #NUM! | 25.7765 | #NUM! | 2.8953 | 0.0014 | 2.7272 | 0.0012 | 7.6347 | #NUM! | 7.5609 | #NUM! | 7.4458 | #NUM! | 102.54 | #NUM! | 101.549 | #NUM! | 1.1611 | 0.1597 | 0.11749 | 9.9313e-05 | 8.978 | 0.002 | 0.8111 | 0.1436 | 0.0262 | 96.4757 | 0.6306 | 2019-07-03 11:21:31 | 434198 | 7.745 | 0 | 7.745 | #NUM! | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.745 | #NUM! | BB
BB-06 | 2019-07-03 | 04:44:18 | 39200053 | 434199 | 49 | 6 | -51 | 49.8 | 10.639 | 0.1 | 5.695 | 0.0148 | 5.7189 | 0.024 | 3.1505 | 0.00095 | 3.152 | 0.001581 | 32.0677 | #NUM! | 32.0617 | #NUM! | 25.2725 | #NUM! | 25.2649 | #NUM! | 2.8519 | 0.0008 | 2.6943 | 0.0015 | 7.1078 | #NUM! | 7.075 | #NUM! | 7.0953 | #NUM! | 100.18 | #NUM! | 99.717 | #NUM! | 0.2693 | 0.0059 | 0.11764 | 0.000205 | 8.941 | 0.002 | 1.1154 | 0.0039 | 0.0015 | 99.9029 | 0.0375 | 2019-07-03 11:26:10 | 434199 | 7.813 | 0.594 | 7.393 | 8.233 | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.813 | 0.84 | BB
BB-06 | 2019-07-03 | 04:45:41 | 39200053 | 434200 | 49 | 6 | -51 | 49.8 | 5.975 | 0.029 | 7.4521 | 0.0007 | 7.4523 | 0.0009 | 3.2722 | 5.9e-05 | 3.2722 | 6.1e-05 | 31.7683 | #NUM! | 31.7686 | #NUM! | 24.8133 | #NUM! | 24.8135 | #NUM! | 2.8004 | 0.0004 | 2.6491 | 0.0007 | 6.651 | #NUM! | 6.6005 | #NUM! | 6.8249 | #NUM! | 97.457 | #NUM! | 96.717 | #NUM! | 0.3429 | 0.0251 | 0.11761 | 0.0001997 | 8.909 | 0 | 0.5222 | 0.0147 | 0.0016 | 99.6338 | 0.0394 | 2019-07-03 11:36:29 | 434200 | 7.855 | 0.001 | 7.854 | 7.856 | 39200053 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.855 | 0.002 | BB
BB-01 | 2019-07-03 | 14:57:27 | 39200058 | 434234 | 48 | 44 | -52 | 58 | 101.593 | 0.076 | -1.103 | 0.0004 | -1.1037 | 0.0007 | 2.6388 | 2.1e-05 | 2.6388 | 3.3e-05 | 32.6071 | #NUM! | 32.6085 | #NUM! | 26.2173 | #NUM! | 26.2185 | #NUM! | 2.4811 | 0.0002 | 2.3304 | 0.0006 | 7.3415 | #NUM! | 7.2486 | #NUM! | 8.3872 | #NUM! | 87.531 | #NUM! | 86.424 | #NUM! | 0.0619 | 0.0069 | 0.11786 | 5.8571e-05 | 8.738 | 0.002 | 3.0341 | -0.0818 | 0.0009 | 102.066 | 0.0215 | 2019-07-03 19:16:21 | 434234 | 8.511 | 0.032 | 8.488 | 8.533 | 39200058 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.5105 | 0.045 | BB
BB-01 | 2019-07-03 | 14:59:50 | 39200058 | 434235 | 48 | 44 | -52 | 58 | 74.453 | 0.107 | -0.559 | 0.0003 | -0.5593 | 0.0002 | 2.6739 | 1.9e-05 | 2.674 | 1.9e-05 | 32.5099 | #NUM! | 32.5108 | #NUM! | 26.1198 | #NUM! | 26.1206 | #NUM! | 2.5885 | 0.0005 | 2.4326 | 0.0006 | 7.5873 | #NUM! | 7.5112 | #NUM! | 8.2722 | #NUM! | 91.72 | #NUM! | 90.801 | #NUM! | 0.0528 | 0.0055 | 0.29235 | 0.0076124 | 8.793 | 0.002 | 3.1037 | -0.074 | 0.0004 | 101.8678 | 0.0101 | 2019-07-03 19:25:59 | 434235 | 8.787 | 0.002 | 8.789 | 8.786 | 39200058 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.7875 | 0.003 | BB
BB-01 | 2019-07-03 | 15:02:28 | 39200058 | 434236 | 48 | 44 | -52 | 58 | 51.048 | 0.087 | 0.0048 | 0.0016 | 0.0062 | 0.0008 | 2.7104 | 0.000118 | 2.7105 | 5.6e-05 | 32.4054 | #NUM! | 32.4056 | #NUM! | 26.0118 | #NUM! | 26.0119 | #NUM! | 2.6549 | 0.0005 | 2.4998 | 0.0008 | 7.6759 | #NUM! | 7.6139 | #NUM! | 8.1561 | #NUM! | 94.113 | #NUM! | 93.352 | #NUM! | 0.0412 | 0.0035 | 1.6044 | 0.018199 | 8.823 | 0.002 | 2.7885 | -0.0568 | 0.0003 | 101.4304 | 0.0077 | 2019-07-03 19:37:51 | 434236 | 8.886 | 0.011 | 8.894 | 8.879 | 39200058 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.8865 | 0.015 | BB
BB-01 | 2019-07-03 | 15:04:04 | 39200058 | 434237 | 48 | 44 | -52 | 58 | 40.65 | 0.065 | 0.5664 | 0.001 | 0.5661 | 0.001 | 2.7483 | 6.8e-05 | 2.7483 | 5.5e-05 | 32.3151 | #NUM! | 32.3161 | #NUM! | 25.9111 | #NUM! | 25.9119 | #NUM! | 2.6878 | 0.0005 | 2.5242 | 0.0003 | 7.6536 | #NUM! | 7.5614 | #NUM! | 8.0425 | #NUM! | 95.165 | #NUM! | 94.019 | #NUM! | 0.1295 | 0.0119 | 3.2073 | 0.025759 | 8.836 | 0.001 | 2.0599 | -0.0546 | 0.0002 | 101.3762 | 0.0054 | 2019-07-03 19:50:39 | 434237 | 8.875 | 0.009 | 8.882 | 8.869 | 39200058 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.8755 | 0.013 | BB
BB-01 | 2019-07-03 | 15:05:31 | 39200058 | 434238 | 48 | 44 | -52 | 58 | 30.706 | 0.227 | 1.1074 | 0.0079 | 1.1109 | 0.0071 | 2.7869 | 0.000479 | 2.7872 | 0.000358 | 32.2547 | #NUM! | 32.254 | #NUM! | 25.8318 | #NUM! | 25.8311 | #NUM! | 2.7396 | 0.0007 | 2.5713 | 0.0006 | 7.7071 | #NUM! | 7.5994 | #NUM! | 7.9342 | #NUM! | 97.138 | #NUM! | 95.781 | #NUM! | 1.5246 | 0.0307 | 6.9243 | 0.19818 | 8.871 | 0 | 2.2798 | 0.0305 | 0.0043 | 99.2407 | 0.1077 | 2019-07-03 20:02:10 | 434238 | 8.96 | 0.04 | 8.932 | 8.988 | 39200058 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.96 | 0.056 | BB
BB-01 | 2019-07-03 | 15:07:00 | 39200058 | 434239 | 48 | 44 | -52 | 58 | 20.825 | 0.366 | 3.7766 | 0.0052 | 3.7742 | 0.0041 | 2.9934 | 0.000243 | 2.9933 | 0.000329 | 32.1036 | #NUM! | 32.1053 | #NUM! | 25.5041 | #NUM! | 25.5056 | #NUM! | 2.8326 | 0.0006 | 2.6693 | 0.0006 | 7.4333 | #NUM! | 7.3632 | #NUM! | 7.4282 | #NUM! | 100.07 | #NUM! | 99.126 | #NUM! | 0.6431 | 0.0048 | 19.112 | 0.86117 | 8.89 | 0.003 | 1.1482 | 0.0109 | 0.0009 | 99.7285 | 0.0216 | 2019-07-03 20:15:13 | 434239 | 8.686 | 0.019 | 8.699 | 8.672 | 39200058 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.6855 | 0.027 | BB
BB-01 | 2019-07-03 | 15:08:25 | 39200058 | 434240 | 48 | 44 | -52 | 58 | 10.583 | 0.031 | 8.345 | 0.04 | 8.353 | 0.036 | 3.3367 | 0.002633 | 3.3371 | 0.002705 | 31.6422 | #NUM! | 31.639 | #NUM! | 24.588 | #NUM! | 24.5843 | #NUM! | 2.8239 | 0.0016 | 2.6738 | 0.0012 | 6.5745 | #NUM! | 6.5335 | #NUM! | 6.6935 | #NUM! | 98.228 | #NUM! | 97.615 | #NUM! | 0.3552 | 0.0078 | 55.284 | 0.47275 | 8.843 | 0.002 | 1.6054 | 0.0345 | 0.0021 | 99.1416 | 0.0532 | 2019-07-03 19:06:34 | 434240 | 7.778 | 0.045 | 7.81 | 7.747 | 39200058 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.7785 | 0.063 | BB
BB-01 | 2019-07-03 | 15:09:21 | 39200058 | 434241 | 48 | 44 | -52 | 58 | 6.304 | 0.086 | 8.5494 | 0.0059 | 8.5521 | 0.0081 | 3.3517 | 0.000423 | 3.3519 | 0.000551 | 31.6173 | #NUM! | 31.6175 | #NUM! | 24.5384 | #NUM! | 24.5382 | #NUM! | 2.8259 | 0.0022 | 2.672 | 0.0007 | 6.534 | #NUM! | 6.5012 | #NUM! | 6.6639 | #NUM! | 98.056 | #NUM! | 97.564 | #NUM! | 0.3203 | 0.0066 | 97.511 | 2.8416 | 8.841 | 0 | 0.9542 | 0.0221 | 0.0011 | 99.4489 | 0.026 | 2019-07-03 18:55:23 | 434241 | 7.713 | 0.017 | 7.725 | 7.701 | 39200058 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.713 | 0.024 | BB
MB-11 | 2019-07-06 | 22:24:03 | 39200079 | 434254 | 55 | 58.7 | -56 | 44.5 | 1199.897 | 0.433 | 3.3653 | 0.0003 | 3.3651 | 0.0002 | 3.2429 | 5.1e-05 | 3.2429 | 4.1e-05 | 34.8904 | #NUM! | 34.8915 | #NUM! | 27.763 | #NUM! | 27.7639 | #NUM! | 2.1239 | 0.0007 | 2.0169 | 0.0004 | 5.9738 | #NUM! | 5.9565 | #NUM! | 7.3629 | #NUM! | 81.136 | #NUM! | 80.9 | #NUM! | 0.0072 | 0.0058 | 0.11771 | 0.00020184 | 8.741 | 0 | 1.9958 | -0.1057 | 0.0002 | 102.6767 | 0.0039 | 2019-07-07 03:18:00 | 434254 | 7.11 | 0.058 | 7.151 | 7.069 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.11 | 0.082 | MB
MB-11 | 2019-07-06 | 22:48:31 | 39200079 | 434255 | 55 | 58.7 | -56 | 44.5 | 152.141 | 0.137 | 3.9781 | 0.0001 | 3.9778 | 0.0001 | 3.2407 | 2.5e-05 | 3.2408 | 2.9e-05 | 34.7653 | #NUM! | 34.7668 | #NUM! | 27.6018 | #NUM! | 27.603 | #NUM! | 2.4406 | 0.0005 | 2.3096 | 0.001 | 6.1382 | #NUM! | 6.1125 | #NUM! | 7.2601 | #NUM! | 84.549 | #NUM! | 84.196 | #NUM! | 0.0197 | 0.0042 | 0.11774 | 0.00019426 | 8.795 | 0.001 | -0.1998 | -0.1078 | 0.0002 | 102.7321 | 0.0039 | 2019-07-07 03:35:15 | 434255 | 7.131 | 0.017 | 7.143 | 7.119 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.131 | 0.024 | MB
MB-11 | 2019-07-06 | 22:51:04 | 39200079 | 434256 | 55 | 58.7 | -56 | 44.5 | 100.608 | 0.253 | 3.771 | 0.0082 | 3.772 | 0.0062 | 3.212 | 0.000977 | 3.2123 | 0.000794 | 34.6707 | #NUM! | 34.673 | #NUM! | 27.5477 | #NUM! | 27.5495 | #NUM! | 2.4865 | 0.001 | 2.3545 | 0.0009 | 6.2712 | #NUM! | 6.2607 | #NUM! | 7.3013 | #NUM! | 85.894 | #NUM! | 85.749 | #NUM! | 0.0365 | 0.0058 | 0.11766 | 0.00020708 | 8.806 | 0 | 0.2464 | -0.1009 | 0.0008 | 102.5549 | 0.0193 | 2019-07-07 03:42:38 | 434256 | 7.308 | 0.003 | 7.306 | 7.31 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.308 | 0.004 | MB
MB-11 | 2019-07-06 | 22:52:48 | 39200079 | 434257 | 55 | 58.7 | -56 | 44.5 | 75.03 | 0.317 | 3.572 | 0.0005 | 3.5718 | 0.0005 | 3.1861 | 7.3e-05 | 3.1862 | 8e-05 | 34.5862 | #NUM! | 34.5878 | #NUM! | 27.5003 | #NUM! | 27.5016 | #NUM! | 2.496 | 0.0005 | 2.3625 | 0.0008 | 6.3315 | #NUM! | 6.2978 | #NUM! | 7.341 | #NUM! | 86.251 | #NUM! | 85.791 | #NUM! | 0.0561 | 0.0091 | 0.11763 | 0.0002036 | 8.813 | 0.002 | 0.3994 | -0.0912 | 0.0008 | 102.3058 | 0.0209 | 2019-07-07 03:54:23 | 434257 | 7.343 | 0.002 | 7.341 | 7.344 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.3425 | 0.003 | MB
MB-11 | 2019-07-06 | 22:54:24 | 39200079 | 434258 | 55 | 58.7 | -56 | 44.5 | 51.606 | 0.065 | 3.9828 | 0.0044 | 3.9807 | 0.0044 | 3.2173 | 0.000156 | 3.2174 | 0.000125 | 34.5364 | #NUM! | 34.5396 | #NUM! | 27.4192 | #NUM! | 27.4219 | #NUM! | 2.57 | 0.0005 | 2.4337 | 0.0006 | 6.4717 | #NUM! | 6.4338 | #NUM! | 7.2705 | #NUM! | 89.016 | #NUM! | 88.494 | #NUM! | 0.41 | 0.0724 | 0.11768 | 0.00020674 | 8.847 | 0.001 | 0.6993 | -0.0484 | 0.0004 | 101.2167 | 0.0101 | 2019-07-07 04:06:17 | 434258 | 7.548 | 0.006 | 7.552 | 7.544 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.548 | 0.008 | MB
MB-11 | 2019-07-06 | 22:56:04 | 39200079 | 434259 | 55 | 58.7 | -56 | 44.5 | 38.989 | 0.211 | 3.99 | 0.0007 | 3.9893 | 0.0007 | 3.1913 | 6.8e-05 | 3.1914 | 0.000107 | 34.2271 | #NUM! | 34.2287 | #NUM! | 27.1724 | #NUM! | 27.1737 | #NUM! | 2.655 | 0.0007 | 2.505 | 0.0008 | 6.7442 | #NUM! | 6.6767 | #NUM! | 7.2844 | #NUM! | 92.587 | #NUM! | 91.66 | #NUM! | 3.5706 | 0.158 | 0.1175 | 0.00012539 | 8.904 | 0.002 | 0.2026 | 0.218 | 0.0033 | 94.6968 | 0.0786 | 2019-07-07 04:17:39 | 434259 | 7.844 | 0.015 | 7.854 | 7.833 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.8435 | 0.021 | MB
MB-11 | 2019-07-06 | 22:57:21 | 39200079 | 434260 | 55 | 58.7 | -56 | 44.5 | 31.22 | 0.319 | 4.3807 | 0.068 | 4.3838 | 0.0743 | 3.2221 | 0.006359 | 3.2224 | 0.006818 | 34.1874 | #NUM! | 34.1881 | #NUM! | 27.0998 | #NUM! | 27.0999 | #NUM! | 2.7309 | 0.0015 | 2.5783 | 0.0019 | 6.9143 | #NUM! | 6.873 | #NUM! | 7.2179 | #NUM! | 95.796 | #NUM! | 95.224 | #NUM! | 5.3901 | 0.2292 | 0.11866 | 0.0020033 | 8.93 | 0 | 0.235 | 0.2898 | 0.0106 | 93.0114 | 0.2466 | 2019-07-07 04:30:13 | 434260 | 8.055 | 0.021 | 8.069 | 8.04 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.0545 | 0.029 | MB
MB-11 | 2019-07-06 | 22:58:26 | 39200079 | 434261 | 55 | 58.7 | -56 | 44.5 | 19.439 | 0.211 | 5.7886 | 0.0067 | 5.7822 | 0.0104 | 3.3304 | 0.000917 | 3.3297 | 0.001569 | 34.0034 | #NUM! | 34.0028 | #NUM! | 26.7915 | #NUM! | 26.7918 | #NUM! | 2.9023 | 0.0009 | 2.7399 | 0.0012 | 7.1462 | #NUM! | 7.0842 | #NUM! | 6.9891 | #NUM! | 102.251 | #NUM! | 101.364 | #NUM! | 3.2115 | 0.1561 | 1.1644 | 0.043669 | 8.987 | 0.001 | 1.5868 | 0.3075 | 0.004 | 92.5995 | 0.0917 | 2019-07-07 04:41:55 | 434261 | 8.336 | 0.006 | 8.34 | 8.332 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.336 | 0.008 | MB
MB-11 | 2019-07-06 | 22:59:54 | 39200079 | 434262 | 55 | 58.7 | -56 | 44.5 | 10.165 | 0.142 | 6.0978 | 0.0019 | 6.0984 | 0.0018 | 3.3556 | 0.000231 | 3.3557 | 0.000206 | 33.9814 | #NUM! | 33.9828 | #NUM! | 26.7354 | #NUM! | 26.7364 | #NUM! | 2.9147 | 0.0007 | 2.7573 | 0.0006 | 7.1211 | #NUM! | 7.0943 | #NUM! | 6.9398 | #NUM! | 102.616 | #NUM! | 102.231 | #NUM! | 2.3755 | 0.1027 | 3.0866 | 0.1774 | 8.999 | 0.001 | 5.7899 | 0.3269 | 0.0027 | 92.1532 | 0.0629 | 2019-07-07 04:53:39 | 434262 | 8.306 | 0.024 | 8.289 | 8.323 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.306 | 0.034 | MB
MB-11 | 2019-07-06 | 23:00:59 | 39200079 | 434263 | 55 | 58.7 | -56 | 44.5 | 4.542 | 0.134 | 6.0701 | 0.0003 | 6.0696 | 0.0004 | 3.3521 | 7.1e-05 | 3.3522 | 7.8e-05 | 33.9737 | #NUM! | 33.9752 | #NUM! | 26.7328 | #NUM! | 26.7341 | #NUM! | 2.9167 | 0.001 | 2.7574 | 0.0022 | 7.1384 | #NUM! | 7.1191 | #NUM! | 6.9446 | #NUM! | 102.794 | #NUM! | 102.517 | #NUM! | 2.4135 | 0.1852 | 9.9883 | 0.52344 | 8.999 | 0 | 60.4476 | 0.3374 | 0.0039 | 91.9113 | 0.09 | 2019-07-07 05:06:17 | 434263 | 8.325 | 0.004 | 8.327 | 8.322 | 39200079 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.3245 | 0.005 | MB
MB-07 | 2019-07-07 | 09:58:53 | 39200083 | 434299 | 55 | 40.7 | -57 | 26.4 | 193.84 | 0.169 | 0.1021 | 0.0032 | 0.1019 | 0.0038 | 2.8135 | 0.00037 | 2.8136 | 0.000428 | 33.5753 | #NUM! | 33.5763 | #NUM! | 26.951 | #NUM! | 26.9518 | #NUM! | 2.3812 | 0.0003 | 2.2527 | 0.0005 | 6.7483 | #NUM! | 6.7207 | #NUM! | 8.0689 | #NUM! | 83.634 | #NUM! | 83.292 | #NUM! | 0.1509 | 0.0151 | 0.11787 | 0 | 8.792 | 0.002 | 0.5984 | -0.0226 | 0.0016 | 100.5662 | 0.0394 | 2019-07-07 15:55:03 | 434299 | 7.533 | 0.363 | 7.276 | 7.79 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.533 | 0.514 | MB
MB-07 | 2019-07-07 | 10:01:31 | 39200083 | 434300 | 55 | 40.7 | -57 | 26.4 | 151.558 | 0.204 | -0.6576 | 0.0017 | -0.6571 | 0.001 | 2.7147 | 0.000153 | 2.7148 | 9.9e-05 | 33.1177 | #NUM! | 33.1185 | #NUM! | 26.6153 | #NUM! | 26.616 | #NUM! | 2.3711 | 0.0007 | 2.2442 | 0.0006 | 6.8633 | #NUM! | 6.8404 | #NUM! | 8.2584 | #NUM! | 83.106 | #NUM! | 82.829 | #NUM! | 0.1521 | 0.0267 | 0.11779 | 0.0001604 | 8.786 | 0.001 | 0.5984 | -0.0451 | 0.001 | 101.1326 | 0.0256 | 2019-07-07 15:43:34 | 434300 | 7.801 | 0.262 | 7.987 | 7.616 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.8015 | 0.371 | MB
MB-07 | 2019-07-07 | 10:04:28 | 39200083 | 434301 | 55 | 40.7 | -57 | 26.4 | 100.058 | 0.308 | -1.1208 | 0.0279 | -1.1392 | 0.0181 | 2.6561 | 0.002661 | 2.6544 | 0.001679 | 32.8631 | #NUM! | 32.8603 | #NUM! | 26.4255 | #NUM! | 26.4238 | #NUM! | 2.354 | 0.0007 | 2.2232 | 0.001 | 6.8556 | #NUM! | 6.798 | #NUM! | 8.376 | #NUM! | 81.848 | #NUM! | 81.16 | #NUM! | 0.1503 | 0.0148 | 0.11778 | 0.00017284 | 8.772 | 0.003 | 0.6012 | -0.0385 | 0 | 100.9681 | 0 | 2019-07-07 15:31:32 | 434301 | 8.158 | 0.298 | 8.369 | 7.947 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.158 | 0.422 | MB
MB-07 | 2019-07-07 | 10:06:44 | 39200083 | 434302 | 55 | 40.7 | -57 | 26.4 | 74.91 | 0.048 | -1.2989 | 0.0028 | -1.2983 | 0.0026 | 2.6319 | 0.000476 | 2.6321 | 0.000478 | 32.7433 | #NUM! | 32.7445 | #NUM! | 26.3336 | #NUM! | 26.3346 | #NUM! | 2.3515 | 0.0009 | 2.2241 | 0.0009 | 6.8683 | #NUM! | 6.8289 | #NUM! | 8.4232 | #NUM! | 81.54 | #NUM! | 81.072 | #NUM! | 0.1908 | 0.0229 | 0.1177 | 0.0002036 | 8.771 | 0 | 1.031 | -0.032 | 0.0004 | 100.802 | 0.009 | 2019-07-07 15:20:30 | 434302 | 7.833 | 0.038 | 7.806 | 7.86 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.833 | 0.054 | MB
MB-07 | 2019-07-07 | 10:08:22 | 39200083 | 434303 | 55 | 40.7 | -57 | 26.4 | 48.516 | 0.092 | -1.4498 | 0.0006 | -1.4495 | 0.0005 | 2.6066 | 1.8e-05 | 2.6067 | 3.4e-05 | 32.5772 | #NUM! | 32.5776 | #NUM! | 26.2029 | #NUM! | 26.2033 | #NUM! | 2.3574 | 0.0005 | 2.2323 | 0.0009 | 6.9055 | #NUM! | 6.8738 | #NUM! | 8.4675 | #NUM! | 81.552 | #NUM! | 81.178 | #NUM! | 0.2704 | 0.0259 | 0.33376 | 0.0074836 | 8.772 | 0 | 0.412 | -0.0278 | 0.0002 | 100.6962 | 0.0039 | 2019-07-07 15:08:53 | 434303 | 8.031 | 0.029 | 8.052 | 8.011 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.0315 | 0.041 | MB
MB-07 | 2019-07-07 | 10:10:03 | 39200083 | 434304 | 55 | 40.7 | -57 | 26.4 | 39.536 | 0.14 | -1.5467 | 0.0017 | -1.5467 | 0.0019 | 2.5949 | 0.000202 | 2.595 | 0.000221 | 32.5281 | #NUM! | 32.5287 | #NUM! | 26.1655 | #NUM! | 26.166 | #NUM! | 2.3611 | 0.0006 | 2.2351 | 0.0003 | 6.9352 | #NUM! | 6.9207 | #NUM! | 8.4926 | #NUM! | 81.661 | #NUM! | 81.49 | #NUM! | 0.1916 | 0.0082 | 0.76779 | 0.015812 | 8.775 | 0.002 | 0.4164 | -0.0246 | 0.0004 | 100.6182 | 0.01 | 2019-07-07 14:56:38 | 434304 | 8.07 | 0.023 | 8.086 | 8.053 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.0695 | 0.033 | MB
MB-07 | 2019-07-07 | 10:11:25 | 39200083 | 434305 | 55 | 40.7 | -57 | 26.4 | 30.206 | 0.101 | -1.5052 | 0.0014 | -1.5053 | 0.0011 | 2.5941 | 8e-05 | 2.5942 | 6.5e-05 | 32.4771 | #NUM! | 32.4778 | #NUM! | 26.123 | #NUM! | 26.1236 | #NUM! | 2.3818 | 0.0008 | 2.2556 | 0.0003 | 6.9918 | #NUM! | 6.9856 | #NUM! | 8.4862 | #NUM! | 82.39 | #NUM! | 82.317 | #NUM! | 0.3861 | 0.0225 | 1.5916 | 0.034068 | 8.783 | 0.002 | 2.2628 | -0.0226 | 0.0004 | 100.5651 | 0.0096 | 2019-07-07 16:05:55 | 434305 | 8.112 | 0.007 | 8.107 | 8.117 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.112 | 0.01 | MB
MB-07 | 2019-07-07 | 10:13:05 | 39200083 | 434306 | 55 | 40.7 | -57 | 26.4 | 20.844 | 0.085 | 0.0079 | 0.019 | 0.0031 | 0.0312 | 2.6847 | 0.001547 | 2.6844 | 0.00238 | 32.0819 | #NUM! | 32.0827 | #NUM! | 25.7505 | #NUM! | 25.7514 | #NUM! | 2.9087 | 0.0018 | 2.7475 | 0.0022 | 8.5377 | #NUM! | 8.489 | #NUM! | 8.1739 | #NUM! | 104.451 | #NUM! | 103.854 | #NUM! | 0.4557 | 0.0299 | 3.4904 | 0.035928 | 9.003 | 0.003 | 3.2073 | -0.0278 | 0.0002 | 100.6968 | 0.0055 | 2019-07-07 16:39:19 | 434306 | 9.988 | 0.025 | 9.97 | 10.005 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 9.9875 | 0.035 | MB
MB-07 | 2019-07-07 | 10:14:59 | 39200083 | 434307 | 55 | 40.7 | -57 | 26.4 | 9.581 | 0.085 | 1.2987 | 0.0072 | 1.2974 | 0.0058 | 2.7421 | 0.000145 | 2.7421 | 0.000112 | 31.5009 | #NUM! | 31.5022 | #NUM! | 25.2149 | #NUM! | 25.2161 | #NUM! | 2.968 | 0.0008 | 2.8095 | 0.0006 | 8.4525 | #NUM! | 8.4385 | #NUM! | 7.9367 | #NUM! | 106.5 | #NUM! | 106.324 | #NUM! | 0.4958 | 0.0407 | 10.115 | 0.30669 | 9.044 | 0 | 3.4441 | -0.0418 | 0 | 101.0495 | 0 | 2019-07-07 16:27:59 | 434307 | 9.78 | 0.016 | 9.791 | 9.769 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 9.78 | 0.022 | MB
MB-07 | 2019-07-07 | 10:16:11 | 39200083 | 434308 | 55 | 40.7 | -57 | 26.4 | 4.947 | 0.083 | 4.2747 | 0.0047 | 4.2751 | 0.0062 | 2.8563 | 0.00018 | 2.8562 | 0.000219 | 30.0329 | #NUM! | 30.0311 | #NUM! | 23.8117 | #NUM! | 23.8102 | #NUM! | 2.8 | 0.0007 | 2.6537 | 0.0006 | 7.3157 | #NUM! | 7.2965 | #NUM! | 7.4418 | #NUM! | 98.308 | #NUM! | 98.05 | #NUM! | 0.44 | 0.0377 | 18.159 | 0.54883 | 9.008 | 0 | 3.727 | 0.0134 | 0.0026 | 99.6648 | 0.0631 | 2019-07-07 16:17:17 | 434308 | 8.461 | 0.044 | 8.492 | 8.43 | 39200083 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.461 | 0.062 | MB
MB-04 | 2019-07-07 | 16:39:27 | 39200086 | 434329 | 55 | 22.3 | -58 | 8.7 | 194.481 | 0.044 | -0.4827 | 0.0134 | -0.4881 | 0.0075 | 2.7512 | 0.00154 | 2.7507 | 0.001136 | 33.39 | #NUM! | 33.3883 | #NUM! | 26.8285 | #NUM! | 26.8274 | #NUM! | 2.346 | 0.0008 | 2.2246 | 0.0004 | 6.763 | #NUM! | 6.7442 | #NUM! | 8.2045 | #NUM! | 82.43 | #NUM! | 82.2 | #NUM! | 0.1509 | 0.0151 | 0.11787 | 0 | 8.785 | 0 | 0.0009 | -0.0368 | 0.0026 | 100.925 | 0.0643 | 2019-07-07 23:59:31 | 434329 | 7.916 | 0.019 | 7.903 | 7.93 | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.9165 | 0.027 | MB
MB-04 | 2019-07-07 | 16:42:22 | 39200086 | 434330 | 55 | 22.3 | -58 | 8.7 | 149.311 | 0.206 | -1.1581 | 0.0004 | -1.1581 | 0.0006 | 2.6529 | 7e-05 | 2.653 | 7.4e-05 | 32.8298 | #NUM! | 32.8316 | #NUM! | 26.3996 | #NUM! | 26.4011 | #NUM! | 2.3348 | 0.0007 | 2.2113 | 0.0005 | 6.8491 | #NUM! | 6.8318 | #NUM! | 8.3863 | #NUM! | 81.669 | #NUM! | 81.464 | #NUM! | 0.1053 | 0.0086 | 0.11782 | 0.00013579 | 8.77 | 0 | 2.6312 | -0.0407 | 0.0003 | 101.023 | 0.0068 | 2019-07-07 23:47:17 | 434330 | 7.952 | 0.013 | 7.962 | 7.943 | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.9525 | 0.019 | MB
MB-04 | 2019-07-07 | 16:45:34 | 39200086 | 434331 | 55 | 22.3 | -58 | 8.7 | 99.965 | 0.064 | -1.4525 | 0.0003 | -1.4527 | 0.0003 | 2.6128 | 4.4e-05 | 2.6129 | 3.5e-05 | 32.6338 | #NUM! | 32.6352 | #NUM! | 26.2489 | #NUM! | 26.2501 | #NUM! | 2.3394 | 0.0007 | 2.2153 | 0.0004 | 6.8886 | #NUM! | 6.8647 | #NUM! | 8.4647 | #NUM! | 81.38 | #NUM! | 81.097 | #NUM! | 0.1365 | 0.0115 | 0.1177 | 0.0002036 | 8.771 | 0.002 | 2.8016 | -0.0299 | 0.0001 | 100.7505 | 0.0039 | 2019-07-07 23:34:22 | 434331 | 7.992 | 0.021 | 8.006 | 7.977 | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.9915 | 0.029 | MB
MB-04 | 2019-07-07 | 16:48:00 | 39200086 | 434332 | 55 | 22.3 | -58 | 8.7 | 75.224 | 0.128 | -1.5539 | 0.0006 | -1.554 | 0.0006 | 2.5943 | 8.3e-05 | 2.5944 | 8.4e-05 | 32.5056 | #NUM! | 32.5067 | #NUM! | 26.1474 | #NUM! | 26.1483 | #NUM! | 2.3585 | 0.0005 | 2.2351 | 0.0003 | 6.9672 | #NUM! | 6.9584 | #NUM! | 8.4956 | #NUM! | 82.009 | #NUM! | 81.904 | #NUM! | 0.2293 | 0.0616 | 0.1178 | 0.00015311 | 8.774 | 0.002 | 2.453 | -0.0246 | 0.0004 | 100.6176 | 0.0093 | 2019-07-07 23:20:35 | 434332 | 8.099 | 0.018 | 8.111 | 8.086 | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.0985 | 0.025 | MB
MB-04 | 2019-07-07 | 16:50:38 | 39200086 | 434333 | 55 | 22.3 | -58 | 8.7 | 50.008 | 0.17 | -1.5255 | 0.0002 | -1.5259 | 0.0002 | 2.5838 | 2e-05 | 2.5839 | 1.8e-05 | 32.3457 | #NUM! | 32.3467 | #NUM! | 26.0168 | #NUM! | 26.0177 | #NUM! | 2.421 | 0.0006 | 2.2915 | 0.0005 | 7.1808 | #NUM! | 7.1537 | #NUM! | 8.4988 | #NUM! | 84.492 | #NUM! | 84.173 | #NUM! | 0.4646 | 0.0219 | 1.1918 | 0.030676 | 8.792 | 0.001 | 1.3571 | -0.0214 | 0.0004 | 100.5368 | 0.01 | 2019-07-07 23:06:14 | 434333 | 8.307 | 0.01 | 8.3 | 8.314 | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.307 | 0.014 | MB
MB-04 | 2019-07-07 | 16:52:44 | 39200086 | 434334 | 55 | 22.3 | -58 | 8.7 | 39.552 | 0.188 | -1.5169 | 0.0002 | -1.5174 | 0.0002 | 2.5832 | 1.8e-05 | 2.5832 | 1.5e-05 | 32.3336 | #NUM! | 32.3348 | #NUM! | 26.0068 | #NUM! | 26.0078 | #NUM! | 2.4332 | 0.0006 | 2.3035 | 0.0003 | 7.2057 | #NUM! | 7.1989 | #NUM! | 8.4975 | #NUM! | 84.796 | #NUM! | 84.717 | #NUM! | 0.42 | 0.0207 | 3.474 | 0.086607 | 8.796 | 0.002 | 1.5161 | -0.0203 | 0.0003 | 100.5086 | 0.0086 | 2019-07-07 22:53:41 | 434334 | 8.406 | 0.016 | 8.395 | 8.418 | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.4065 | 0.023 | MB
MB-04 | 2019-07-07 | 16:54:39 | 39200086 | 434335 | 55 | 22.3 | -58 | 8.7 | 30.651 | 0.114 | -1.296 | 0.0067 | -1.3003 | 0.0048 | 2.5937 | 0.000432 | 2.5935 | 0.000376 | 32.2438 | #NUM! | 32.2468 | #NUM! | 25.9284 | #NUM! | 25.9309 | #NUM! | 2.517 | 0.0005 | 2.3835 | 0.0006 | 7.4672 | #NUM! | 7.4541 | #NUM! | 8.4525 | #NUM! | 88.342 | #NUM! | 88.187 | #NUM! | 0.7679 | 0.0346 | 9.5416 | 0.19419 | 8.84 | 0.002 | 1.2102 | -0.0212 | 0.0003 | 100.5323 | 0.0066 | 2019-07-07 22:41:28 | 434335 | 8.738 | 0.009 | 8.731 | 8.744 | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.7375 | 0.013 | MB
MB-04 | 2019-07-07 | 16:56:31 | 39200086 | 434336 | 55 | 22.3 | -58 | 8.7 | 20.692 | 0.159 | 0.2934 | 0.0202 | 0.2969 | 0.0113 | 2.707 | 0.001923 | 2.7075 | 0.001302 | 32.0758 | #NUM! | 32.0776 | #NUM! | 25.7323 | #NUM! | 25.7336 | #NUM! | 2.9036 | 0.0007 | 2.7331 | 0.0007 | 8.4651 | #NUM! | 8.3924 | #NUM! | 8.1134 | #NUM! | 104.334 | #NUM! | 103.439 | #NUM! | 1.5085 | 0.0965 | 33.824 | 0.80226 | 9.017 | 0 | 1.3982 | 0.0047 | 0.0137 | 99.8824 | 0.3417 | 2019-07-07 22:33:10 | 434336 | 9.811 | 0 | 9.811 | #NUM! | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 9.811 | #NUM! | MB
MB-04 | 2019-07-07 | 16:58:53 | 39200086 | 434337 | 55 | 22.3 | -58 | 8.7 | 10.579 | 0.064 | 1.9803 | 0.0016 | 1.9778 | 0.0014 | 2.793 | 5.2e-05 | 2.7931 | 0.000139 | 31.4566 | #NUM! | 31.4607 | #NUM! | 25.1351 | #NUM! | 25.1385 | #NUM! | 3.0416 | 0.001 | 2.8624 | 0.0007 | 8.5537 | #NUM! | 8.4809 | #NUM! | 7.8023 | #NUM! | 109.632 | #NUM! | 108.699 | #NUM! | 0.3254 | 0.0095 | 111.25 | 1.5451 | 9.082 | 0 | 0.5463 | 0.0084 | 0.0018 | 99.7905 | 0.0443 | 2019-07-07 22:13:18 | 434337 | 9.937 | 0.007 | 9.942 | 9.932 | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 9.937 | 0.01 | MB
MB-04 | 2019-07-07 | 17:00:35 | 39200086 | 434338 | 55 | 22.3 | -58 | 8.7 | 5.248 | 0.083 | 2.811 | 0.0052 | 2.8042 | 0.0041 | 2.7407 | 0.000788 | 2.742 | 0.000765 | 30.018 | #NUM! | 30.0395 | #NUM! | 23.9249 | #NUM! | 23.9425 | #NUM! | 2.9119 | 0.0006 | 2.7461 | 0.0006 | 7.9911 | #NUM! | 7.9362 | #NUM! | 7.7165 | #NUM! | 103.561 | #NUM! | 102.849 | #NUM! | 0.324 | 0.0104 | 171.87 | 7.087 | 9.057 | 0 | 5.052 | 0.0663 | 0.0084 | 98.3554 | 0.2082 | 2019-07-07 22:00:10 | 434338 | 9.51 | 0.252 | 9.418 | 9.318 | 39200086 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 9.368 | 0.1 | MB
SI-04 | 2019-07-08 | 19:51:01 | 39200101 | 434388 | 53 | 32 | -55 | 8.7 | 200.275 | 0.075 | 0.5368 | 0.0019 | 0.5359 | 0.0008 | 2.87 | 0.000191 | 2.87 | 8.7e-05 | 33.8356 | #NUM! | 33.8362 | #NUM! | 27.1373 | #NUM! | 27.1378 | #NUM! | 2.3753 | 0.0005 | 2.2431 | 0 | 6.6304 | #NUM! | 6.5835 | #NUM! | 7.9638 | #NUM! | 83.258 | #NUM! | 82.668 | #NUM! | 0.1108 | 0.011 | 0.11783 | 0.00012539 | 8.769 | 0.001 | 2.7863 | -0.0362 | 0.0011 | 100.9083 | 0.0266 | 2019-07-09 01:50:02 | 434388 | 7.677 | 0 | 7.677 | #NUM! | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.677 | #NUM! | SI
SI-04 | 2019-07-08 | 19:53:44 | 39200101 | 434389 | 53 | 32 | -55 | 8.7 | 150.932 | 0.137 | 0.1666 | 0.0013 | 0.1671 | 0.0004 | 2.8254 | 0.000152 | 2.8255 | 4.6e-05 | 33.6862 | #NUM! | 33.687 | #NUM! | 27.0371 | #NUM! | 27.0377 | #NUM! | 2.41 | 0.0007 | 2.2739 | 0.0004 | 6.7955 | #NUM! | 6.7441 | #NUM! | 8.0491 | #NUM! | 84.425 | #NUM! | 83.786 | #NUM! | 0.1239 | 0.0116 | 0.1178 | 0.00015311 | 8.791 | 0.001 | 0.1675 | -0.0641 | 0.0002 | 101.6175 | 0.0066 | 2019-07-09 01:25:25 | 434389 | 7.834 | 0.008 | 7.839 | 7.828 | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.8335 | 0.011 | SI
SI-04 | 2019-07-08 | 19:56:28 | 39200101 | 434390 | 53 | 32 | -55 | 8.7 | 98.502 | 0.137 | -1.0041 | 0.009 | -1.0074 | 0.0024 | 2.688 | 0.001036 | 2.6876 | 0.000277 | 33.1692 | #NUM! | 33.1683 | #NUM! | 26.6697 | #NUM! | 26.6691 | #NUM! | 2.4085 | 0.0006 | 2.2726 | 0.0004 | 7.0189 | #NUM! | 6.9666 | #NUM! | 8.3318 | #NUM! | 84.241 | #NUM! | 83.614 | #NUM! | 0.0899 | 0.0094 | 0.11768 | 0.00020674 | 8.792 | 0.002 | 0.6889 | -0.0485 | 0.0005 | 101.2189 | 0.0118 | 2019-07-09 00:43:08 | 434390 | 8.598 | 0.734 | 9.117 | 8.079 | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.598 | 1.038 | SI
SI-04 | 2019-07-08 | 19:58:44 | 39200101 | 434391 | 53 | 32 | -55 | 8.7 | 75.87 | 0.045 | -1.1755 | 0.0006 | -1.1764 | 0.0008 | 2.6608 | 3e-05 | 2.6607 | 4.9e-05 | 33.0012 | #NUM! | 33.0021 | #NUM! | 26.5391 | #NUM! | 26.5399 | #NUM! | 2.4271 | 0.0006 | 2.287 | 0.0006 | 7.1132 | #NUM! | 7.0494 | #NUM! | 8.3801 | #NUM! | 84.882 | #NUM! | 84.121 | #NUM! | 0.2898 | 0.0423 | 0.11817 | 0.00059744 | 8.794 | 0.002 | 1.0573 | -0.0381 | 0.0005 | 100.9565 | 0.0136 | 2019-07-09 01:00:54 | 434391 | 8.196 | 0.017 | 8.184 | 8.208 | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.196 | 0.024 | SI
SI-04 | 2019-07-08 | 20:00:40 | 39200101 | 434392 | 53 | 32 | -55 | 8.7 | 49.759 | 0.212 | -1.0986 | 0.0046 | -1.0998 | 0.0061 | 2.6523 | 0.000267 | 2.6522 | 0.000296 | 32.817 | #NUM! | 32.8175 | #NUM! | 26.3874 | #NUM! | 26.3878 | #NUM! | 2.5243 | 0.0005 | 2.3752 | 0.0007 | 7.4345 | #NUM! | 7.3537 | #NUM! | 8.3737 | #NUM! | 88.783 | #NUM! | 87.818 | #NUM! | 1.7152 | 0.0637 | 1.0105 | 0.036194 | 8.828 | 0 | 1.9196 | 0.014 | 0.0086 | 99.6515 | 0.2141 | 2019-07-09 00:49:11 | 434392 | 8.574 | 0.013 | 8.565 | 8.583 | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.574 | 0.018 | SI
SI-04 | 2019-07-08 | 20:02:17 | 39200101 | 434393 | 53 | 32 | -55 | 8.7 | 39.87 | 0.077 | -1.0629 | 0.0136 | -1.03 | 0.0263 | 2.644 | 0.001136 | 2.6465 | 0.002042 | 32.6713 | #NUM! | 32.6697 | #NUM! | 26.2681 | #NUM! | 26.2658 | #NUM! | 2.6913 | 0.0016 | 2.5258 | 0.0024 | 8.0658 | #NUM! | 8.004 | #NUM! | 8.3744 | #NUM! | 96.315 | #NUM! | 95.576 | #NUM! | 0.4114 | 0.0302 | 2.2193 | 0.010441 | 8.879 | 0.002 | 1.8127 | -0.018 | 0.014 | 100.4527 | 0.351 | 2019-07-09 00:37:08 | 434393 | 9.172 | 0 | 9.172 | #NUM! | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 9.172 | #NUM! | SI
SI-04 | 2019-07-08 | 20:04:09 | 39200101 | 434394 | 53 | 32 | -55 | 8.7 | 29.914 | 0.264 | -0.7926 | 0.0288 | -0.7649 | 0.0436 | 2.6571 | 0.002138 | 2.6593 | 0.003293 | 32.5615 | #NUM! | 32.5609 | #NUM! | 26.1702 | #NUM! | 26.1687 | #NUM! | 2.8993 | 0.0015 | 2.7118 | 0.0014 | 8.7468 | #NUM! | 8.6275 | #NUM! | 8.3207 | #NUM! | 105.121 | #NUM! | 103.688 | #NUM! | 0.3391 | 0.023 | 4.1839 | 0.10969 | 8.957 | 0.001 | 0.5195 | -0.028 | 0.0005 | 100.7012 | 0.0115 | 2019-07-09 02:35:12 | 434394 | 9.995 | 0.059 | 10.037 | 9.954 | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 9.9955 | 0.083 | SI
SI-04 | 2019-07-08 | 20:05:41 | 39200101 | 434395 | 53 | 32 | -55 | 8.7 | 19.98 | 0.254 | 1.6801 | 0.0077 | 1.6861 | 0.0082 | 2.8233 | 0.00104 | 2.8241 | 0.000577 | 32.1315 | #NUM! | 32.1363 | #NUM! | 25.6963 | #NUM! | 25.6997 | #NUM! | 3.1054 | 0.001 | 2.8988 | 0.0006 | 8.8159 | #NUM! | 8.6582 | #NUM! | 7.8255 | #NUM! | 112.657 | #NUM! | 110.641 | #NUM! | 0.5713 | 0.0449 | 7.6629 | 0.127 | 8.985 | 0.002 | 1.4618 | -0.0208 | 0.0007 | 100.5224 | 0.0166 | 2019-07-09 02:21:55 | 434395 | 10.253 | 0.015 | 10.263 | 10.242 | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 10.2525 | 0.021 | SI
SI-04 | 2019-07-08 | 20:07:10 | 39200101 | 434396 | 53 | 32 | -55 | 8.7 | 11.037 | 0.079 | 4.8673 | 0.0078 | 4.8676 | 0.0098 | 2.9832 | 0.000502 | 2.9832 | 0.000619 | 30.9458 | #NUM! | 30.9462 | #NUM! | 24.4765 | #NUM! | 24.4768 | #NUM! | 2.8664 | 0.001 | 2.6841 | 0.0008 | 7.353 | #NUM! | 7.2258 | #NUM! | 7.291 | #NUM! | 100.853 | #NUM! | 99.108 | #NUM! | 0.4169 | 0.0205 | 14.817 | 0.089402 | 9.001 | 0.002 | 2.3516 | 0.0323 | 0.008 | 99.1969 | 0.1972 | 2019-07-09 02:09:00 | 434396 | 8.485 | 0.045 | 8.453 | 8.517 | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.485 | 0.064 | SI
SI-04 | 2019-07-08 | 20:08:31 | 39200101 | 434397 | 53 | 32 | -55 | 8.7 | 5.797 | 0.045 | 5.4859 | 0.0021 | 5.4858 | 0.0023 | 3.0199 | 0.000139 | 3.0199 | 0.000159 | 30.7932 | #NUM! | 30.7936 | #NUM! | 24.2889 | #NUM! | 24.2892 | #NUM! | 2.8387 | 0.0006 | 2.6605 | 0.0007 | 7.155 | #NUM! | 7.0312 | #NUM! | 7.1913 | #NUM! | 99.499 | #NUM! | 97.776 | #NUM! | 0.3493 | 0.041 | 31.104 | 0.28198 | 8.994 | 0 | 0.3714 | 0.025 | 0.0019 | 99.378 | 0.0482 | 2019-07-09 01:56:47 | 434397 | 8.313 | 0.001 | 8.314 | 8.312 | 39200101 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.313 | 0.002 | SI
SI-09 | 2019-07-09 | 06:09:11 | 39200106 | 434430 | 54 | 12 | -54 | 0 | 211.809 | 0.131 | 2.7748 | 0.0003 | 2.7749 | 0.0001 | 3.1104 | 2.3e-05 | 3.1105 | 1.8e-05 | 34.4422 | #NUM! | 34.4428 | #NUM! | 27.4601 | #NUM! | 27.4605 | #NUM! | 2.4106 | 0.0003 | 2.2778 | 0.0005 | 6.3137 | #NUM! | 6.2552 | #NUM! | 7.4936 | #NUM! | 84.256 | #NUM! | 83.476 | #NUM! | 0.1685 | 0.0081 | 0.11786 | 5.8571e-05 | 8.778 | 0.002 | 11.4222 | -0.012 | 0.0012 | 100.2988 | 0.031 | 2019-07-09 10:23:18 | 434430 | 7.351 | 0.011 | 7.359 | 7.344 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.3515 | 0.015 | SI
SI-09 | 2019-07-09 | 06:12:16 | 39200106 | 434431 | 54 | 12 | -54 | 0 | 149.258 | 0.233 | -0.5329 | 0.0006 | -0.5329 | 0.0007 | 2.7405 | 5.1e-05 | 2.7406 | 6.2e-05 | 33.3286 | #NUM! | 33.3294 | #NUM! | 26.7809 | #NUM! | 26.7816 | #NUM! | 2.384 | 0.0004 | 2.2477 | 0.0006 | 6.8641 | #NUM! | 6.7967 | #NUM! | 8.2189 | #NUM! | 83.516 | #NUM! | 82.695 | #NUM! | 0.06 | 0.0085 | 0.11768 | 0.00020674 | 8.79 | 0.001 | 0.0151 | -0.0543 | 0.0005 | 101.3657 | 0.0142 | 2019-07-09 10:35:00 | 434431 | 7.908 | 0.006 | 7.904 | 7.912 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.908 | 0.008 | SI
SI-09 | 2019-07-09 | 06:14:48 | 39200106 | 434432 | 54 | 12 | -54 | 0 | 99.535 | 0.112 | -0.9819 | 0.0023 | -0.9868 | 0.007 | 2.6779 | 0.000342 | 2.6773 | 0.000896 | 33.0067 | #NUM! | 33.005 | #NUM! | 26.5373 | #NUM! | 26.5361 | #NUM! | 2.3728 | 0.0005 | 2.2312 | 0.0004 | 6.8969 | #NUM! | 6.8139 | #NUM! | 8.3365 | #NUM! | 82.731 | #NUM! | 81.735 | #NUM! | 0.1105 | 0.0079 | 0.11781 | 0.00014496 | 8.779 | 0 | 0.8961 | -0.0514 | 0.0002 | 101.2931 | 0.0039 | 2019-07-09 10:46:41 | 434432 | 7.932 | 0.018 | 7.945 | 7.919 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.932 | 0.026 | SI
SI-09 | 2019-07-09 | 06:16:44 | 39200106 | 434433 | 54 | 12 | -54 | 0 | 75.816 | 0.236 | -1.1544 | 0.0002 | -1.1543 | 0.0001 | 2.6516 | 2.5e-05 | 2.6516 | 2.9e-05 | 32.8532 | #NUM! | 32.8537 | #NUM! | 26.4185 | #NUM! | 26.4188 | #NUM! | 2.3678 | 0.0007 | 2.2206 | 0.0006 | 6.9039 | #NUM! | 6.799 | #NUM! | 8.3841 | #NUM! | 82.344 | #NUM! | 81.094 | #NUM! | 0.1318 | 0.0297 | 0.11776 | 0.00018289 | 8.77 | 0.001 | 0.2859 | -0.0413 | 0.0005 | 101.0379 | 0.0136 | 2019-07-09 10:59:54 | 434433 | 7.926 | 0.012 | 7.918 | 7.935 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.9265 | 0.017 | SI
SI-09 | 2019-07-09 | 06:18:49 | 39200106 | 434434 | 54 | 12 | -54 | 0 | 50.053 | 0.204 | -1.2353 | 0.0002 | -1.2353 | 0.0002 | 2.6365 | 1.9e-05 | 2.6365 | 2.3e-05 | 32.7507 | #NUM! | 32.751 | #NUM! | 26.3378 | #NUM! | 26.3381 | #NUM! | 2.3722 | 0.0006 | 2.229 | 0.0004 | 6.9117 | #NUM! | 6.8299 | #NUM! | 8.4084 | #NUM! | 82.199 | #NUM! | 81.226 | #NUM! | 0.1244 | 0.013 | 0.11784 | 0.00011342 | 8.771 | 0 | 0.6105 | -0.0333 | 0.0003 | 100.8347 | 0.0075 | 2019-07-09 11:12:01 | 434434 | 7.991 | 0.021 | 7.976 | 8.005 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.9905 | 0.029 | SI
SI-09 | 2019-07-09 | 06:20:21 | 39200106 | 434435 | 54 | 12 | -54 | 0 | 39.812 | 0.092 | -1.1789 | 0.0028 | -1.1773 | 0.0015 | 2.638 | 0.000131 | 2.6381 | 8.1e-05 | 32.7163 | #NUM! | 32.716 | #NUM! | 26.3082 | #NUM! | 26.3079 | #NUM! | 2.395 | 0.0005 | 2.2544 | 0.0005 | 6.9775 | #NUM! | 6.9107 | #NUM! | 8.3978 | #NUM! | 83.087 | #NUM! | 82.292 | #NUM! | 0.579 | 0.0641 | 0.1178 | 0.00015311 | 8.776 | 0.002 | 2.1454 | -0.029 | 0.0019 | 100.7289 | 0.0485 | 2019-07-09 11:25:00 | 434435 | 8.03 | 0.006 | 8.034 | 8.025 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.0295 | 0.009 | SI
SI-09 | 2019-07-09 | 06:21:53 | 39200106 | 434436 | 54 | 12 | -54 | 0 | 30.591 | 0.327 | -1.2683 | 0.025 | -1.271 | 0.0271 | 2.6228 | 0.001578 | 2.6226 | 0.001632 | 32.6115 | #NUM! | 32.6122 | #NUM! | 26.2258 | #NUM! | 26.2265 | #NUM! | 2.3965 | 0.0006 | 2.2545 | 0 | 7.004 | #NUM! | 6.9217 | #NUM! | 8.4242 | #NUM! | 83.141 | #NUM! | 82.164 | #NUM! | 1.1643 | 0.2266 | 0.11776 | 0.00018289 | 8.771 | 0 | 1.9908 | -0.0273 | 0.0006 | 100.6835 | 0.0137 | 2019-07-09 11:37:56 | 434436 | 8.12 | 0.052 | 8.157 | 8.083 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.12 | 0.074 | SI
SI-09 | 2019-07-09 | 06:23:25 | 39200106 | 434437 | 54 | 12 | -54 | 0 | 19.911 | 0.236 | -1.4564 | 0.0045 | -1.4595 | 0.0108 | 2.6006 | 0.000236 | 2.6007 | 0.000693 | 32.5198 | #NUM! | 32.524 | #NUM! | 26.1565 | #NUM! | 26.16 | #NUM! | 2.4561 | 0.0013 | 2.3017 | 0.0018 | 7.2736 | #NUM! | 7.1028 | #NUM! | 8.4725 | #NUM! | 85.849 | #NUM! | 83.834 | #NUM! | 3.2507 | 0.3315 | 0.11757 | 0.00018289 | 8.801 | 0.001 | 1.8001 | 0.0985 | 0.0093 | 97.5681 | 0.2275 | 2019-07-09 11:50:37 | 434437 | 8.418 | 0.098 | 8.488 | 8.349 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.4185 | 0.139 | SI
SI-09 | 2019-07-09 | 06:24:57 | 39200106 | 434438 | 54 | 12 | -54 | 0 | 10.155 | 0.054 | 2.2408 | 0.0079 | 2.2427 | 0.0149 | 2.8459 | 0.000492 | 2.846 | 0.00094 | 31.8512 | #NUM! | 31.8511 | #NUM! | 25.4321 | #NUM! | 25.4319 | #NUM! | 2.9024 | 0.0006 | 2.7061 | 0.0008 | 7.9906 | #NUM! | 7.8375 | #NUM! | 7.7302 | #NUM! | 103.37 | #NUM! | 101.39 | #NUM! | 0.8412 | 0.1664 | 0.11764 | 0.000205 | 8.995 | 0.001 | 2.5868 | 0.0213 | 0.01 | 99.4694 | 0.2483 | 2019-07-09 12:03:20 | 434438 | 9.216 | 0.001 | 9.217 | 9.216 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 9.2165 | 0.001 | SI
SI-09 | 2019-07-09 | 06:26:22 | 39200106 | 434439 | 54 | 12 | -54 | 0 | 5.505 | 0.062 | 2.6192 | 0.0028 | 2.6156 | 0.0008 | 2.8712 | 0.000199 | 2.871 | 7.4e-05 | 31.7872 | #NUM! | 31.788 | #NUM! | 25.3521 | #NUM! | 25.353 | #NUM! | 2.8995 | 0.0005 | 2.7097 | 0.0008 | 7.8955 | #NUM! | 7.7639 | #NUM! | 7.6605 | #NUM! | 103.07 | #NUM! | 101.351 | #NUM! | 0.5329 | 0.0703 | 0.11771 | 0.00020184 | 8.997 | 0.002 | 2.5588 | -0.0008 | 0.0044 | 100.0191 | 0.1106 | 2019-07-09 12:16:37 | 434439 | 9.154 | 0.03 | 9.176 | 9.133 | 39200106 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 9.1545 | 0.043 | SI
SI-13 | 2019-07-09 | 16:20:24 | 39200110 | 434462 | 54 | 47 | -53 | 0 | 1075.189 | 0.499 | 3.3816 | 0.0002 | 3.382 | 0.0002 | 3.2372 | 3.8e-05 | 3.2373 | 3.9e-05 | 34.8695 | #NUM! | 34.8696 | #NUM! | 27.7448 | #NUM! | 27.7449 | #NUM! | 2.1986 | 0.0006 | 2.0826 | 0.0005 | 6.1519 | #NUM! | 6.1188 | #NUM! | 7.361 | #NUM! | 83.576 | #NUM! | 83.126 | #NUM! | 0.0132 | 0.0052 | 0.11786 | 5.8571e-05 | 8.751 | 0 | 3.0894 | -0.1074 | 0.0005 | 102.721 | 0.0133 | 2019-07-09 21:31:07 | 434462 | 7.289 | 0.008 | 7.294 | 7.283 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.2885 | 0.011 | SI
SI-13 | 2019-07-09 | 16:36:09 | 39200110 | 434463 | 54 | 47 | -53 | 0 | 500.245 | 0.058 | 3.427 | 0.0007 | 3.4274 | 0.0006 | 3.2126 | 6.1e-05 | 3.2126 | 8.5e-05 | 34.8275 | #NUM! | 34.8276 | #NUM! | 27.7069 | #NUM! | 27.7069 | #NUM! | 2.3835 | 0.0006 | 2.2537 | 0.0006 | 6.3304 | #NUM! | 6.2802 | #NUM! | 7.3549 | #NUM! | 86.072 | #NUM! | 85.389 | #NUM! | 0.0113 | 0.0037 | 0.11769 | 0.000205 | 8.787 | 0.002 | -0.5906 | -0.1057 | 0 | 102.6773 | 0 | 2019-07-09 21:43:08 | 434463 | 7.357 | 0.002 | 7.356 | 7.359 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.3575 | 0.003 | SI
SI-13 | 2019-07-09 | 16:43:25 | 39200110 | 434464 | 54 | 47 | -53 | 0 | 250.819 | 0.138 | 3.572 | 0.0002 | 3.5724 | 0.0002 | 3.2122 | 1.7e-05 | 3.2122 | 1.6e-05 | 34.8028 | #NUM! | 34.8032 | #NUM! | 27.6729 | #NUM! | 27.6732 | #NUM! | 2.4391 | 0.0005 | 2.3057 | 0.0005 | 6.2836 | #NUM! | 6.2298 | #NUM! | 7.3302 | #NUM! | 85.725 | #NUM! | 84.99 | #NUM! | 0.0185 | 0.0037 | 0.11769 | 0.00020604 | 8.798 | 0.002 | -0.4804 | -0.1056 | 0.0003 | 102.6756 | 0.0066 | 2019-07-09 21:55:02 | 434464 | 7.293 | 0.015 | 7.303 | 7.282 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.2925 | 0.021 | SI
SI-13 | 2019-07-09 | 16:46:40 | 39200110 | 434465 | 54 | 47 | -53 | 0 | 150.008 | 0.355 | 3.7197 | 0.0005 | 3.7202 | 0.0015 | 3.2196 | 5.5e-05 | 3.2197 | 0.000117 | 34.79 | #NUM! | 34.7902 | #NUM! | 27.6479 | #NUM! | 27.6481 | #NUM! | 2.4575 | 0.001 | 2.3243 | 0.0007 | 6.2237 | #NUM! | 6.1825 | #NUM! | 7.3045 | #NUM! | 85.205 | #NUM! | 84.641 | #NUM! | 0.0255 | 0.0083 | 0.11774 | 0.00019092 | 8.801 | 0 | 2.192 | -0.1066 | 0.0003 | 102.7022 | 0.0075 | 2019-07-09 22:06:31 | 434465 | 7.223 | 0.009 | 7.229 | 7.216 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.2225 | 0.013 | SI
SI-13 | 2019-07-09 | 16:49:11 | 39200110 | 434466 | 54 | 47 | -53 | 0 | 99.617 | 0.329 | 3.8044 | 0.002 | 3.8101 | 0.0012 | 3.2241 | 0.000149 | 3.2246 | 9.3e-05 | 34.7808 | #NUM! | 34.7803 | #NUM! | 27.632 | #NUM! | 27.631 | #NUM! | 2.4678 | 0.0005 | 2.3352 | 0.0004 | 6.2122 | #NUM! | 6.1676 | #NUM! | 7.2899 | #NUM! | 85.218 | #NUM! | 84.607 | #NUM! | 0.0359 | 0.0052 | 0.11768 | 0.00020674 | 8.804 | 0.002 | 0.0211 | -0.1078 | 0 | 102.7315 | 0 | 2019-07-09 22:18:03 | 434466 | 7.203 | 0.025 | 7.221 | 7.185 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.203 | 0.036 | SI
SI-13 | 2019-07-09 | 16:51:29 | 39200110 | 434467 | 54 | 47 | -53 | 0 | 76.305 | 0.281 | 3.823 | 0.001 | 3.8241 | 0.0007 | 3.2222 | 2.9e-05 | 3.2223 | 2.1e-05 | 34.7513 | #NUM! | 34.7512 | #NUM! | 27.6066 | #NUM! | 27.6064 | #NUM! | 2.4797 | 0.0005 | 2.3476 | 0.0012 | 6.2311 | #NUM! | 6.1746 | #NUM! | 7.2881 | #NUM! | 85.499 | #NUM! | 84.724 | #NUM! | 0.0582 | 0.0052 | 0.34354 | 0.0058772 | 8.805 | 0 | 0.1165 | -0.1016 | 0.0008 | 102.5715 | 0.0209 | 2019-07-09 22:29:23 | 434467 | 7.196 | 0.001 | 7.197 | 7.195 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.196 | 0.002 | SI
SI-13 | 2019-07-09 | 16:53:48 | 39200110 | 434468 | 54 | 47 | -53 | 0 | 48.927 | 0.162 | 3.9456 | 0.0002 | 3.9458 | 0.0004 | 3.2257 | 3e-05 | 3.2258 | 3.6e-05 | 34.6779 | #NUM! | 34.6785 | #NUM! | 27.5356 | #NUM! | 27.5361 | #NUM! | 2.5462 | 0.0008 | 2.4056 | 0.0013 | 6.39 | #NUM! | 6.3663 | #NUM! | 7.2701 | #NUM! | 87.897 | #NUM! | 87.57 | #NUM! | 0.1448 | 0.0086 | 1.3483 | 0.015311 | 8.826 | 0.002 | 0.9898 | -0.0833 | 0.0003 | 102.1058 | 0.0086 | 2019-07-09 23:41:32 | 434468 | 7.471 | 0.013 | 7.462 | 7.48 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.471 | 0.018 | SI
SI-13 | 2019-07-09 | 16:55:31 | 39200110 | 434469 | 54 | 47 | -53 | 0 | 39.119 | 0.456 | 4.0237 | 0.002 | 4.0237 | 0.0023 | 3.2263 | 4.5e-05 | 3.2264 | 3.6e-05 | 34.6076 | #NUM! | 34.6081 | #NUM! | 27.4715 | #NUM! | 27.4719 | #NUM! | 2.5764 | 0.0004 | 2.4311 | 0.0006 | 6.4713 | #NUM! | 6.4045 | #NUM! | 7.2598 | #NUM! | 89.141 | #NUM! | 88.221 | #NUM! | 0.2656 | 0.0086 | 3.1947 | 0.12431 | 8.842 | 0.001 | 0.986 | -0.0656 | 0.0013 | 101.653 | 0.0321 | 2019-07-09 23:28:46 | 434469 | 7.549 | 0.056 | 7.509 | 7.588 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.5485 | 0.079 | SI
SI-13 | 2019-07-09 | 16:58:22 | 39200110 | 434470 | 54 | 47 | -53 | 0 | 29.874 | 0.469 | 4.0804 | 0.0027 | 4.0788 | 0.0039 | 3.2246 | 0.00056 | 3.2247 | 0.000152 | 34.5318 | #NUM! | 34.5349 | #NUM! | 27.4054 | #NUM! | 27.408 | #NUM! | 2.601 | 0.0007 | 2.4562 | 0.0005 | 6.5283 | #NUM! | 6.4792 | #NUM! | 7.2535 | #NUM! | 90.004 | #NUM! | 89.326 | #NUM! | 0.6826 | 0.0358 | 5.3145 | 0.092266 | 8.86 | 0.002 | 2.9688 | -0.0295 | 0.0013 | 100.7406 | 0.0327 | 2019-07-09 23:16:30 | 434470 | 7.614 | 0.006 | 7.618 | 7.61 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.614 | 0.008 | SI
SI-13 | 2019-07-09 | 17:00:09 | 39200110 | 434471 | 54 | 47 | -53 | 0 | 20.325 | 0.163 | 4.602 | 0.0135 | 4.6028 | 0.0116 | 3.2622 | 0.000952 | 3.2622 | 0.000852 | 34.4345 | #NUM! | 34.433 | #NUM! | 27.2718 | #NUM! | 27.2706 | #NUM! | 2.6814 | 0.0006 | 2.526 | 0.0012 | 6.6825 | #NUM! | 6.5956 | #NUM! | 7.1678 | #NUM! | 93.233 | #NUM! | 92.019 | #NUM! | 4.0554 | 0.28 | 13.5 | 0.17384 | 8.896 | 0.001 | 1.7957 | 0.1501 | 0.0078 | 96.3179 | 0.1893 | 2019-07-09 23:04:06 | 434471 | 7.7 | 0.134 | 7.795 | 7.605 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 7.7 | 0.19 | SI
SI-13 | 2019-07-09 | 17:02:09 | 39200110 | 434472 | 54 | 47 | -53 | 0 | 9.509 | 0.226 | 6.424 | 0.0209 | 6.4237 | 0.0204 | 3.3727 | 0.003978 | 3.3727 | 0.003863 | 33.8475 | #NUM! | 33.8479 | #NUM! | 26.5878 | #NUM! | 26.5881 | #NUM! | 2.9468 | 0.0014 | 2.783 | 0.0009 | 7.1539 | #NUM! | 7.1029 | #NUM! | 6.8934 | #NUM! | 103.783 | #NUM! | 103.042 | #NUM! | 1.9832 | 0.0819 | 225.8 | 16.761 | 8.997 | 0.001 | 1.7008 | 0.5145 | 0.0099 | 87.9316 | 0.2188 | 2019-07-09 22:52:25 | 434472 | 8.421 | 0.018 | 8.433 | 8.408 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.4205 | 0.025 | SI
SI-13 | 2019-07-09 | 17:04:00 | 39200110 | 434473 | 54 | 47 | -53 | 0 | 4.688 | 0.081 | 5.8859 | 0.0013 | 5.8856 | 0.0009 | 3.2191 | 0.000377 | 3.2194 | 0.00036 | 32.6586 | #NUM! | 32.663 | #NUM! | 25.7167 | #NUM! | 25.7202 | #NUM! | 2.8749 | 0.0006 | 2.7129 | 0.0004 | 7.1043 | #NUM! | 7.0434 | #NUM! | 7.0358 | #NUM! | 100.977 | #NUM! | 100.112 | #NUM! | 0.4224 | 0.008 | 349.93 | 22.223 | 9.003 | 0.002 | 28.4099 | 0.1963 | 0.0043 | 95.2117 | 0.1023 | 2019-07-09 22:41:01 | 434473 | 8.277 | 0.01 | 8.27 | 8.284 | 39200110 | 1221 | 3313 | 0.4339 | 3539 | 0.4637 | 8.277 | 0.014 | SI

## tel200_meanO2.txt
Analysis_Date_time | Sample | [O2](ml/l) | std_[O2](ml/l) | Rep_1(ml/l) | Rep_2(ml/l)
2019-06-27 15:05:10 | 433878 | 6.717 | 0.042 | 6.747 | 6.688
2019-06-27 15:17:07 | 433879 | 8.695 | 0.047 | 8.662 | 8.729
2019-06-27 15:29:57 | 433880 | 8.847 | 0.124 | 8.759 | 8.935
2019-06-27 16:15:06 | 433881 | 8.803 | 0.008 | 8.798 | 8.809
2019-06-27 16:26:32 | 433882 | 8.726 | 0.006 | 8.721 | 8.73
2019-06-27 16:37:00 | 433883 | 8.422 | 0.021 | 8.407 | 8.436
2019-06-27 16:47:26 | 433884 | 8.173 | 0.003 | 8.175 | 8.171
2019-06-27 16:57:20 | 433885 | 8.063 | 0.047 | 8.097 | 8.03
2019-06-27 17:06:51 | 433886 | 7.962 | 0.011 | 7.969 | 7.954
2019-06-27 22:51:44 | 433907 | 6.225 | 0.011 | 6.233 | 6.217
2019-06-27 23:02:41 | 433914 | 7.769 | 0 | 7.769 | #NUM!
2019-06-27 23:08:34 | 433913 | 7.78 | 0.006 | 7.775 | 7.784
2019-06-27 23:24:23 | 433908 | 7.806 | 0.033 | 7.783 | 7.829
2019-06-27 23:36:31 | 433909 | 8.605 | 0.066 | 8.652 | 8.558
2019-06-27 23:47:15 | 433910 | 8.695 | 0.017 | 8.683 | 8.707
2019-06-27 23:59:11 | 433911 | 8.149 | 0.012 | 8.157 | 8.14
2019-06-28 00:09:41 | 433912 | 7.956 | 0.004 | 7.959 | 7.953
2019-06-28 22:07:50 | 433958 | 6.16 | 0.029 | 6.181 | 6.14
2019-06-28 22:19:36 | 433959 | 7.525 | 0.001 | 7.526 | 7.524
2019-06-28 22:30:35 | 433960 | 8.352 | 0.001 | 8.353 | 8.351
2019-06-28 22:40:54 | 433961 | 8.465 | 0.007 | 8.46 | 8.47
2019-06-28 22:52:07 | 433962 | 8.708 | 0.013 | 8.717 | 8.699
2019-06-28 23:03:31 | 433963 | 8.873 | 0.009 | 8.88 | 8.867
2019-06-28 23:14:48 | 433964 | 8.591 | 0.03 | 8.612 | 8.57
2019-06-28 23:26:02 | 433965 | 8.242 | 0.015 | 8.253 | 8.232
2019-06-28 23:36:39 | 433966 | 8.029 | 0.018 | 8.042 | 8.017
2019-06-28 23:47:57 | 433967 | 8.042 | 0.01 | 8.049 | 8.035
2019-06-29 14:19:19 | 434010 | 7.454 | 0.008 | 7.46 | 7.449
2019-06-29 14:31:18 | 434009 | 7.47 | 0.008 | 7.475 | 7.464
2019-06-29 14:42:40 | 434008 | 7.755 | 0.021 | 7.77 | 7.74
2019-06-29 14:54:13 | 434007 | 7.878 | 0.001 | 7.878 | 7.879
2019-06-29 15:04:24 | 434006 | 8.06 | 0.013 | 8.069 | 8.051
2019-06-29 15:15:25 | 434005 | 7.666 | 0.028 | 7.686 | 7.646
2019-06-29 15:42:14 | 434004 | 6.739 | 0.004 | 6.736 | 6.742
2019-06-29 15:55:52 | 434003 | 6.629 | 0 | 6.629 | 6.629
2019-06-29 16:06:18 | 434002 | 6.303 | 0.009 | 6.296 | 6.309
2019-06-29 16:21:34 | 434001 | 6.499 | 0.011 | 6.507 | 6.491
2019-06-30 00:03:52 | 434037 | 7.587 | 0.018 | 7.6 | 7.574
2019-06-30 00:14:44 | 434036 | 7.354 | 0.006 | 7.358 | 7.35
2019-06-30 00:25:08 | 434035 | 6.418 | 0.004 | 6.415 | 6.42
2019-06-30 00:33:47 | 434034 | 6.328 | 0.011 | 6.32 | 6.336
2019-06-30 00:44:27 | 434033 | 5.804 | 0.041 | 5.775 | 5.833
2019-06-30 00:52:58 | 434032 | 5.523 | 0.006 | 5.518 | 5.527
2019-06-30 01:01:00 | 434038 | 7.853 | 0.002 | 7.854 | 7.851
2019-06-30 01:12:39 | 434039 | 7.936 | 0.008 | 7.93 | 7.942
2019-06-30 01:24:54 | 434040 | 7.921 | 0.016 | 7.909 | 7.932
2019-06-30 01:37:14 | 434041 | 7.529 | 0.007 | 7.524 | 7.534
2019-06-30 15:10:42 | 434075 | 7.053 | 0.001 | 7.054 | 7.053
2019-06-30 15:22:44 | 434076 | 7.131 | 0.001 | 7.132 | 7.13
2019-06-30 15:33:49 | 434077 | 6.87 | 0.005 | 6.874 | 6.867
2019-06-30 15:44:16 | 434078 | 6.122 | 0.006 | 6.117 | 6.126
2019-06-30 15:53:23 | 434079 | 6.08 | 0.008 | 6.074 | 6.086
2019-06-30 16:04:17 | 434080 | 5.985 | 0.022 | 6 | 5.969
2019-06-30 16:15:12 | 434086 | 7.832 | 0.034 | 7.856 | 7.808
2019-06-30 16:32:43 | 434084 | 8.248 | 0.032 | 8.226 | 8.271
2019-06-30 16:43:56 | 434083 | 7.048 | 0.013 | 7.038 | 7.057
2019-06-30 16:54:06 | 434082 | 6.778 | 0.011 | 6.786 | 6.771
2019-06-30 17:04:47 | 434081 | 6.476 | 0.009 | 6.47 | 6.483
2019-07-02 05:30:57 | 434105 | 7.523 | 0.035 | 7.548 | 7.498
2019-07-02 05:42:43 | 434106 | 7.622 | 0.077 | 7.677 | 7.568
2019-07-02 05:56:25 | 434107 | 7.724 | 0.022 | 7.74 | 7.709
2019-07-02 06:08:09 | 434108 | 8.791 | 0.001 | 8.792 | 8.79
2019-07-02 06:19:14 | 434110 | 8.675 | 0.024 | 8.658 | 8.692
2019-07-02 06:30:33 | 434099 | 7.23 | 0.014 | 7.24 | 7.22
2019-07-02 06:40:45 | 434100 | 7.285 | 0.052 | 7.248 | 7.322
2019-07-02 06:51:03 | 434101 | 7.299 | 0.031 | 7.321 | 7.277
2019-07-02 07:01:55 | 434102 | 7.405 | 0.039 | 7.432 | 7.377
2019-07-02 07:14:37 | 434103 | 7.488 | 0.015 | 7.498 | 7.477
2019-07-02 07:27:40 | 434104 | 7.32 | 0.008 | 7.314 | 7.326
2019-07-02 22:51:39 | 434156 | 8.424 | 0.034 | 8.448 | 8.4
2019-07-02 23:02:27 | 434155 | 8.429 | 0.02 | 8.443 | 8.415
2019-07-02 23:13:24 | 434154 | 8.442 | 0.004 | 8.444 | 8.439
2019-07-02 23:24:29 | 434153 | 8.541 | 0.007 | 8.536 | 8.546
2019-07-02 23:34:51 | 434147 | 7.043 | 0.002 | 7.045 | 7.042
2019-07-02 23:44:19 | 434148 | 7.531 | 0.003 | 7.529 | 7.533
2019-07-02 23:54:31 | 434149 | 7.778 | 0.004 | 7.775 | 7.781
2019-07-03 00:04:16 | 434150 | 7.995 | 0.03 | 8.016 | 7.974
2019-07-03 00:14:29 | 434151 | 8.748 | 0.028 | 8.729 | 8.768
2019-07-03 00:24:51 | 434152 | 8.535 | 0.015 | 8.525 | 8.546
2019-07-03 09:42:23 | 434196 | 8.687 | 0.016 | 8.698 | 8.676
2019-07-03 10:06:54 | 434195 | 8.548 | 0 | 8.548 | #NUM!
2019-07-03 10:12:56 | 434194 | 8.344 | 0.033 | 8.32 | 8.367
2019-07-03 10:28:34 | 434193 | 8.231 | 0.103 | 8.304 | 8.158
2019-07-03 10:43:17 | 434192 | 7.931 | 0.032 | 7.909 | 7.954
2019-07-03 10:53:18 | 434191 | 6.651 | 0.003 | 6.653 | 6.649
2019-07-03 11:03:31 | 434197 | 8.311 | 0.675 | 7.834 | 8.789
2019-07-03 11:21:31 | 434198 | 7.745 | 0 | 7.745 | #NUM!
2019-07-03 11:26:10 | 434199 | 7.813 | 0.594 | 7.393 | 8.233
2019-07-03 11:36:29 | 434200 | 7.855 | 0.001 | 7.854 | 7.856
2019-07-03 18:55:23 | 434241 | 7.713 | 0.017 | 7.725 | 7.701
2019-07-03 19:06:34 | 434240 | 7.778 | 0.045 | 7.81 | 7.747
2019-07-03 19:16:21 | 434234 | 8.511 | 0.032 | 8.488 | 8.533
2019-07-03 19:25:59 | 434235 | 8.787 | 0.002 | 8.789 | 8.786
2019-07-03 19:37:51 | 434236 | 8.886 | 0.011 | 8.894 | 8.879
2019-07-03 19:50:39 | 434237 | 8.875 | 0.009 | 8.882 | 8.869
2019-07-03 20:02:10 | 434238 | 8.96 | 0.04 | 8.932 | 8.988
2019-07-03 20:15:13 | 434239 | 8.686 | 0.019 | 8.699 | 8.672
2019-07-07 03:18:00 | 434254 | 7.11 | 0.058 | 7.151 | 7.069
2019-07-07 03:35:15 | 434255 | 7.131 | 0.017 | 7.143 | 7.119
2019-07-07 03:42:38 | 434256 | 7.308 | 0.003 | 7.306 | 7.31
2019-07-07 03:54:23 | 434257 | 7.343 | 0.002 | 7.341 | 7.344
2019-07-07 04:06:17 | 434258 | 7.548 | 0.006 | 7.552 | 7.544
2019-07-07 04:17:39 | 434259 | 7.844 | 0.015 | 7.854 | 7.833
2019-07-07 04:30:13 | 434260 | 8.055 | 0.021 | 8.069 | 8.04
2019-07-07 04:41:55 | 434261 | 8.336 | 0.006 | 8.34 | 8.332
2019-07-07 04:53:39 | 434262 | 8.306 | 0.024 | 8.289 | 8.323
2019-07-07 05:06:17 | 434263 | 8.325 | 0.004 | 8.327 | 8.322
2019-07-07 14:56:38 | 434304 | 8.07 | 0.023 | 8.086 | 8.053
2019-07-07 15:08:53 | 434303 | 8.031 | 0.029 | 8.052 | 8.011
2019-07-07 15:20:30 | 434302 | 7.833 | 0.038 | 7.806 | 7.86
2019-07-07 15:31:32 | 434301 | 8.158 | 0.298 | 8.369 | 7.947
2019-07-07 15:43:34 | 434300 | 7.801 | 0.262 | 7.987 | 7.616
2019-07-07 15:55:03 | 434299 | 7.533 | 0.363 | 7.276 | 7.79
2019-07-07 16:05:55 | 434305 | 8.112 | 0.007 | 8.107 | 8.117
2019-07-07 16:17:17 | 434308 | 8.461 | 0.044 | 8.492 | 8.43
2019-07-07 16:27:59 | 434307 | 9.78 | 0.016 | 9.791 | 9.769
2019-07-07 16:39:19 | 434306 | 9.988 | 0.025 | 9.97 | 10.005
2019-07-07 22:00:10 | 434338 | 9.51 | 0.252 | 9.418 | 9.318
2019-07-07 22:13:18 | 434337 | 9.937 | 0.007 | 9.942 | 9.932
2019-07-07 22:33:10 | 434336 | 9.811 | 0 | 9.811 | #NUM!
2019-07-07 22:41:28 | 434335 | 8.738 | 0.009 | 8.731 | 8.744
2019-07-07 22:53:41 | 434334 | 8.406 | 0.016 | 8.395 | 8.418
2019-07-07 23:06:14 | 434333 | 8.307 | 0.01 | 8.3 | 8.314
2019-07-07 23:20:35 | 434332 | 8.099 | 0.018 | 8.111 | 8.086
2019-07-07 23:34:22 | 434331 | 7.992 | 0.021 | 8.006 | 7.977
2019-07-07 23:47:17 | 434330 | 7.952 | 0.013 | 7.962 | 7.943
2019-07-07 23:59:31 | 434329 | 7.916 | 0.019 | 7.903 | 7.93
2019-07-09 00:37:08 | 434393 | 9.172 | 0 | 9.172 | #NUM!
2019-07-09 00:43:08 | 434390 | 8.598 | 0.734 | 9.117 | 8.079
2019-07-09 00:49:11 | 434392 | 8.574 | 0.013 | 8.565 | 8.583
2019-07-09 01:00:54 | 434391 | 8.196 | 0.017 | 8.184 | 8.208
2019-07-09 01:25:25 | 434389 | 7.834 | 0.008 | 7.839 | 7.828
2019-07-09 01:50:02 | 434388 | 7.677 | 0 | 7.677 | #NUM!
2019-07-09 01:56:47 | 434397 | 8.313 | 0.001 | 8.314 | 8.312
2019-07-09 02:09:00 | 434396 | 8.485 | 0.045 | 8.453 | 8.517
2019-07-09 02:21:55 | 434395 | 10.253 | 0.015 | 10.263 | 10.242
2019-07-09 02:35:12 | 434394 | 9.995 | 0.059 | 10.037 | 9.954
2019-07-09 10:23:18 | 434430 | 7.351 | 0.011 | 7.359 | 7.344
2019-07-09 10:35:00 | 434431 | 7.908 | 0.006 | 7.904 | 7.912
2019-07-09 10:46:41 | 434432 | 7.932 | 0.018 | 7.945 | 7.919
2019-07-09 10:59:54 | 434433 | 7.926 | 0.012 | 7.918 | 7.935
2019-07-09 11:12:01 | 434434 | 7.991 | 0.021 | 7.976 | 8.005
2019-07-09 11:25:00 | 434435 | 8.03 | 0.006 | 8.034 | 8.025
2019-07-09 11:37:56 | 434436 | 8.12 | 0.052 | 8.157 | 8.083
2019-07-09 11:50:37 | 434437 | 8.418 | 0.098 | 8.488 | 8.349
2019-07-09 12:03:20 | 434438 | 9.216 | 0.001 | 9.217 | 9.216
2019-07-09 12:16:37 | 434439 | 9.154 | 0.03 | 9.176 | 9.133
2019-07-09 21:31:07 | 434462 | 7.289 | 0.008 | 7.294 | 7.283
2019-07-09 21:43:08 | 434463 | 7.357 | 0.002 | 7.356 | 7.359
2019-07-09 21:55:02 | 434464 | 7.293 | 0.015 | 7.303 | 7.282
2019-07-09 22:06:31 | 434465 | 7.223 | 0.009 | 7.229 | 7.216
2019-07-09 22:18:03 | 434466 | 7.203 | 0.025 | 7.221 | 7.185
2019-07-09 22:29:23 | 434467 | 7.196 | 0.001 | 7.197 | 7.195
2019-07-09 22:41:01 | 434473 | 8.277 | 0.01 | 8.27 | 8.284
2019-07-09 22:52:25 | 434472 | 8.421 | 0.018 | 8.433 | 8.408
2019-07-09 23:04:06 | 434471 | 7.7 | 0.134 | 7.795 | 7.605
2019-07-09 23:16:30 | 434470 | 7.614 | 0.006 | 7.618 | 7.61
2019-07-09 23:28:46 | 434469 | 7.549 | 0.056 | 7.509 | 7.588
2019-07-09 23:41:32 | 434468 | 7.471 | 0.013 | 7.462 | 7.48

## failed titrations
Sample | Station-ID | Sbeox0ML/L | WINKLER | WINKLER-DIFF
434197 | BB-06 | 7.6229 | 8.3115 | 0.955
434199 | BB-06 | 7.1078 | 7.813 | 0.84
434299 | MB-07 | 6.7483 | 7.533 | 0.514
434300 | MB-07 | 6.8633 | 7.8015 | 0.371
434301 | MB-07 | 6.8556 | 8.158 | 0.422
434306 | MB-07 | 8.5377 | 9.9875 | 0.035
434307 | MB-07 | 8.4525 | 9.78 | 0.022
434336 | MB-04 | 8.4651 | 9.811 | #NUM!
434337 | MB-04 | 8.5537 | 9.937 | 0.01
434338 | MB-04 | 7.9911 | 9.368 | 0.1
434390 | SI-04 | 7.0189 | 8.598 | 1.038
434393 | SI-04 | 8.0658 | 9.172 | #NUM!
434394 | SI-04 | 8.7468 | 9.9955 | 0.083
434395 | SI-04 | 8.8159 | 10.2525 | 0.021
434438 | SI-09 | 7.9906 | 9.2165 | 0.001
434439 | SI-09 | 7.8955 | 9.1545 | 0.043